# Hall A Tritium Target System

R. Holt, C. Keith, D. Meekins, T. O’Connor et.al.

Abstract

The following is a working document describing the design, construction, installation, operation, and removal of the Hall a tritium target system currently under development at Jefferson Lab. The current target design is a result of efforts from numerous collaborators from various institutions. It must be understood that this is a working document. It shall be updated as the design of the target system evolves and as new lessons are learned. Numerous supporting calculations and technical reports are referenced throughout the document and may be found on the tritium target wiki.

Contents

1	Definitions and Acronyms	3

2	Introduction	5

2.1	General description	5

2.2	Operating summary	6

3	System Overview	8

3.1	Subsystems	8

4	Tritium Containment	11

4.1	Shipping	11

4.2	Installation and Removal	11

4.3	Operating Mode	12

5	Target Motion and Cryostat	13

5.1	Motion System	13

5.2	Cryostat	13

6	Target Cell	16

6.1	Design	16

6.1.1	Materials	18

6.1.2	Static Pressure Design	18

6.1.3	Thermal and Cyclic Design	18

6.1.4	Filling and Shipping	23

6.1.5	Prototype cells	26

6.1.6	Other Gas Cells	27

6.2	Tritium Permeation	27

6.3	Tritium Embrittlement and Corrosion	28

7	Cell Assembly and Installation/Removal	30

7.1	Cleaning	30

7.2	Inspection	30

7.3	Assembly	30

7.4	Testing	31

7.5	Cell Purification	31

7.6	Installation and Removal	31

8	Vacuum System	33

8.1	Design	34

8.2	Getter System	35

8.3	Collimator and Isolation Window	37

8.4	Alarms	39

9	Control System	40

9.1	Epics Controls for the JLAB Tritium Target	40

9.2	Overview of Target Operation	40

9.3	Control Hardware and Interlocks	41

9.3.1	Machine Fast Shutdown Interlocks	43

10	Shipping	45

11	Target Cell Activation	46

12	Exhaust System	47

12.1	Stack and Blower	47

12.2	Vacuum Exhaust	47

12.3	Makeup Air	49

13	Handling Hut	50

13.1	Hut Design	50

13.2	Installation and Removal	50

14	Tritium Detection and Monitoring	51

15	Cell Handling	52

15.1	Filling/Recovery	52

15.2	Shipping	52

15.2.1	Secondary Containment Vessel	52

15.2.2	Outer Shipping Vessel	52

15.3	Installation and Removal	52

16	Release of Tritium	54

16.1	Case 1	54

16.2	Case 2	54

16.3	Case 3	55

16.3.1	Case 3 A)	55

16.3.2	Case 3 B)	55

16.3.3	Case 3 C)	55

16.4	Effects of Release to Environment Outside of Hall A	56

16.4.1	Assumptions Made in the Models	56

17	Operational Performance	58

17.1	Tritium Content	58

18	Bio-Assay Program	60

19	NEPA	61

20	Procedures	62

20.1	Hall Access Procedures	62

20.1.1	Truck Ramp Protocols	63

20.1.2	Controlled Access	63

20.1.3	Restricted Access	64

20.1.4	Special Access for Installation and Removal of Cell	64

21	Hazard Analysis	66

21.1	Pressure hazards	67

21.2	Radiological Hazards	67

22	Response to Committee (June 3, 2010) Report	69

22.1	Administrative Action Items	69

22.1.1	Task 1:	69

22.1.2	Task 2:	69

22.1.3	Task 3:	69

22.1.4	Task 4:	69

22.1.5	Task 5:	70

22.1.6	Task 6:	70

22.1.7	Task 7:	70

22.2	Target Cell	70

22.2.1	Task 1:	70

22.2.2	Task 2:	70

22.2.3	Task 3:	70

22.2.4	Task 4:	71

22.2.5	Task 5:	71

22.2.6	Task 6:	71

22.2.7	Task 7:	71

22.2.8	Task 8:	71

22.2.9	Task 9:	72

22.2.10	Task 10:	72

22.2.11	Task 11:	72

22.2.12	Task 12:	72

22.2.13	Task 13:	72

22.2.14	Task 14:	72

22.2.15	Task 15:	72

22.2.16	Task 16:	72

22.2.17	Task 17:	73

22.2.18	Task 18:	73

22.2.19	Task 19:	73

22.2.20	Task 20	73

22.3	73

22.3.1	Task 1	73

22.3.2	Task 2	73

22.3.3	Task 3	74

22.3.4	Task 4	74

22.3.5	Task 5	74

22.3.6	Task 6	74

22.3.7	Task 7	74

22.3.8	Task 8	74

22.3.9	Task 9	74

22.3.10	Task 10	74

22.4	EHS Tasks	74

22.4.1	Task 1	75

22.4.2	Task 2	75

22.4.3	Task 3	75

22.4.4	Task 4	76

22.4.5	Task 5	76

22.4.6	Task 6	76

22.4.7	Task 7	76

22.4.8	Task 8	76

23	Summary of Calculations	77

# Definitions and Acronyms

ASCE:  American Society of Civil Engineers

ASCE-7:  Minimum Design Loads for Buildings and Other Structures

ASME: American Society of Mechanical Engineers

ASME STS-1: ASME Standard for Steel Stacks

B31.3:  ASME B31.3 Process Piping Code

BPVC:  ASME Boiler and Pressure Vessel Code

CFD: Computational Fluid Dynamics

DAC:  Derived Air Concentration

EDM:  Extensible Display Manager for EPICS

EPICS:   Experimental Physics and Industrial Control System

FEA: Finite Element Analysis

FMEA: Failure Modes and Effects Analysis

FMECA: Failure Modes Effects and Criticality Analysis

FML:  JLAB Facilities Management and Logistic group.

FS:  Factor of Safety

ICO:  Interagency Cooperative Agreement

MEI:  Maximally Exposed Individual

P&ID: Process and Instrumentation Diagram

PID:  Proportional-integral-derivative control-feedback loop

RWP:  radiological work permit

SMC:  Stepper Motor Controller

SRS:  Savannah River Site

SRNL:  Savannah River National Laboratory

SRTE: Savannah River Tritium Enterprises

TED:  Total Effective Dose

Beam Permit:  No one is allowed in this area and a safety sweep has been conducted to confirm that the area has been vacated. All access doors are locked and safety interlocks are engaged. The accelerator is powered and the beam can be turned on at any time.

Controlled Access: The accelerator is powered but there is no beam present in the Hall. A limited number of people following strict safety procedures may enter this are to perform specific tasks or make specific repairs.

Restricted Access:  The Hall is open for general access for any person with proper training. This condition shall be modified for tritium target operations.

Shall:  Indicates a required action, construction, or design element.

Should:  Indicates that an action is considered best (or at least better) practice.

# Introduction

## General description

The JLAB Hall A Tritium Target is expected to be used for a number of approved experiments (see Refs [1–4]) starting in the Fall of 2016 at Jefferson Lab Continuous Electron Beam Accelerator Facility (CEBAF). The target consists of 6 cells of various species of gas installed in place of the standard Hall A cryogenic target. The Qweak target motion system and cryostat shall be altered to be compatible with the tritium target. Several solid targets will also be hung below the cells to form a target ladder (see Figure 1). Each of the cells will be filled to the required pressure and valved and capped off at room temperature. The cells are thus considered to be sealed. The gaseous fluid will not be circulated marking a departure from the normal Hall A target operation.

The intent is to design and construct a simple and robust target system which mitigates the hazards associated with tritium by minimizing the amount of tritium required to perform the experiments and by multiple layers of containment. While shipping and while at JLAB the tritium will be enclosed by three layers of containment. These layers are detailed in in Section 4.

The target cells are of a modular design. This allows for easy assembly of the target system and permits installation of the tritium cell after all other target installation activities have been completed. The modular design chosen also allows the tritium cell to be filled off site at Savannah River site tritium Enterprises (SRTE). This limits the amount of tritium at Jefferson Lab to the amount contained in the cell which is slightly less than 1100 Ci. In effect, Jefferson Lab will not “handle” tritium.

The entire target assembly shall be cooled with 15K helium from the End Station Refrigerator (ESR). The 15K helium will be preheated to 40K and used to cool a heat sink attached to the cell assemblies. This will remove the ~15W of heat that is generated by the electron beam. The beam current allowed on the tritium cell shall also be limited to a maximum of 20 µA.

Figure 1:  Target ladder with gas cells and solid targets. Note the position of the tritium cell is at the top of the ladder.

## Operating summary

The filled cell is expected to be at Jefferson Lab for less than one calendar year. Upon completion of the experiments, the cell shall be shipped back to SRTE where the tritium will be recovered. The tritium cell will contain 1099 Ci tritium gas which at room temperature gives a pressure in the cell of about 200 psia.

The target, once installed, shall operate in three modes:

- Warm or standby mode at room temperature

  - Standby mode for no beam operations

- Cool down/warm up mode with temperatures between 40K and 300K

  - This is a temporary mode while the target is prepared for beam or warmed for standby.

- Run mode with temperature at 40K

  - This is normal run mode for beam on target

The target cells will be filled to various pressures depending on the fill gas as indicated in the table below:

| Target | Fill Gas/Solid material | Pressure at (300K) | Max Beam Current (µA) |
| --- | --- | --- | --- |
| Tritium | Tritium (T2) | 200 psia | 20 |
| Hydrogen | 2H | 500 psia | 20 |
| Deuterium | 2D | 500 psia | 20 |
| Helium | 3He | 500 psia | 25 |
| Argon | Ar | 500 psia | 25 |
| Empty | N/A | 0 psia | 25 |
| Alignment target | Al 7075 | N/A | 20 |
| Carbon optics | C | N/A | 25 |
| Carbon Hole | C | N/A | 25 |
| Carbon | C | N/A | 25 |
| BeO Viewer | BeO | N/A | 25 |
| Solid 1 | TBD | N/A | TBD |
| Solid 2 | TBD | N/A | TBD |

Table 1: Target ladder list and maximum currents.

The cells will be attached to a heat sink cooled by the End Station Refrigerator (ESR) and maintained at 50K by a heater controlled by a PID.

# System Overview

The target is designed using a systems approach which incorporates a number of engineering and administrative controls to ensure a safety and performance. The target system must fully support beam operations for six gas cells and a number of solid targets.

## Subsystems

The Hall A Tritium Target System shall contain the following subsystems:

- Target system local

  - Motion system

  - Instrumentation and Control system

  - Cells and solid targets

  - Cryogenic system

- Scattering Chamber

  - Vacuum system

  - Isolation system (Be window)

  - Getter system

- Exhaust System

  - Hall A high velocity ventilation

  - Vacuum pump exhaust

  - 20m exterior stack (20m above grade at site boundary)

- Tritium detection/monitoring system

  - Portable monitors

  - Fixed remote monitors

- Administrative/Procedural

  - Required to maintain multiple layers of containment.

Note that there are no substantive changes required for the beamline. Thus, with the exception of the raster, it will not be considered as part of the target system.

The above subsystems are discussed in detail in the following sections. The local system and scattering chamber are shown in the figures below.

Figure 2:  Internal local target system. This figure depicts the target ladder (gas cells and solid target), the motion system and the cryostat.

Figure 3:  Full assembly view of the local target vacuum systems. Note that the target is in the “HOME” position and that the top cell (tritium) is not visible. It is protected by the thick walls of the chamber in this position.

# Tritium Containment

Tritium shall be contained by at least 3 layers of containment at all times during shipping, insulation and removal, and operating phases. Thus, while at JLAB and while being transported there shall be triple layer of containment for the tritium. It shall be assumed that while the cell is at SRTE, their policies and procedures provide a suitable level of safety. The 3 layers of containment can only be assured through proper application of administrative, procedural, and engineering controls. These are discussed in the sections below as they apply to each phase of the tritium target lifecycle.

Some of the necessary administrative and procedural controls required are:

- Restricting access to Hall a via the CANS (electronic access) system.

  - Limit access to trained personnel (both users and staff).

- Parking the target in the home position during restricted and controlled access.

- Controlling the truck ramp doors both top and bottom using CANS system.

- Detailed procedures for installation and removal of the tritium cell from the scattering chamber.

- Detailed procedures for taking the Hall from being permit to controlled access.

- Detailed procedures for taking the Hall from controlled access to restricted access.

The purpose of the containment system is to 1) prevent any release of tritium and 2) to control any release such that exposers to the public and personnel are minimized. See Section 16 for more details regarding the release of tritium.

## Shipping

During shipping cycle, the cell shall be considered the primary containment for the tritium. A second intermediate container provides the next layer of containment and allows for sampling of this intermediate volume for tritium leakage. The 2nd vessel shall be fully capable of containing the full volume of tritium in the cell. This vessel shall also have custom formed foam padding that will protect the tritium cell assembly during shipping. The 3rd level of containment shall be a DOT approved pressure vessel acceptable for Type A shipments. While a great deal of progress has been made on this issue, the shipping system is still under development at this time.

## Installation and Removal

During the installation and removal stages, the primary containment of the tritium shall be the tritium cell. Secondary containment level shall be affected by the handling hut. The 3rd level of containment shall be provided by Hall A itself. For Hall A to provide containment, it shall be placed in controlled access and all closed to all non-essential personnel. All ramp doors shall remain closed. Personnel access doors may be opened to allow access but shall not be propped open. While the handling hut and Hall A are considered levels of containment in this model, should the primary containment fail, all tritium in the cell shall be collected by the exhaust system and stacked to the environment. This exhaust system is discussed in Section 12. The release scenario is discussed in Section 16. Failure to follow procedures could result in a worst-case incident whereby the entire contents of the tritium cell would be released the Hall A. Tritium could then exit the Hall through the truck ramp and expose workers and the public to tritium. Note that no injury or overexposure to workers or the public is expected in such a worst-case incident.

## Operating Mode

The system shall be considered to be an operating mode when the cell is installed on the target ladder in the scattering chamber is fully sealed. In this case, the cell is the primary containment and the scattering chamber is a secondary containment. The Hall can be in one of 3 conditions during the operating mode. These conditions are:

- Restricted access:  under this condition beam is not available to the hall and personnel they enter the Hall without RadCon escort.

- Controlled access: under this condition beam is ready to be delivered to the hall but the Hall is being accessed with a RadCon escort. This is a temporary condition intended for short-term maintenance activities.

- Beam permit: under this condition personnel are not allowed in the hall and beam may be delivered.

While in controlled access and beam permit, all doors including the truck ramp rollup doors must remain closed. Under these 2 conditions the Hall may be considered a 3rd level of containment. However, should a breach of the primary and or secondary containment occur, the tritium would be released to the environment through the exhaust system. A breach of the primary containment system (target cell) would be captured by the vacuum system (see Section 8).

Restricted access presents a number of challenges. This access condition cannot be banned for practical reasons. Therefore, in order for the Hall to function as a third level of containment, administrative and procedural controls must be in place and followed. These include:

- Badge access to the Hall must be activated through the CANS system.

- This includes operation of the truck ramp door at the bottom of the ramp.

- Access through the truck ramp doors must be performed as though passing through an air lock system. The bottom door may not be opened until the top ramp door is closed and similarly the top door may not be opened if the bottom door is not closed.

- Installation of a louvered door at the personnel access portal at the base of the ramp.

- Installation of the exhaust system to extract contaminated air from the all to the stack.

- Installation of scattering chamber window covers prior to entering restricted access.

- Installation of manual activation for the exhaust system.

- Use of portable and fixed (remote) tritium monitors.

  - The portable monitor shall give an audible alarm indicating that the exhaust system must be manually activated and the Hall evacuated.

  - The remote monitor will also give an alarm but shall activate the exhaust system automatically.

- All personnel in the Hall must have specific tritium training.

By following these procedures and enacting appropriate engineering controls the Hall may be considered a containment level with regard to preventing an uncontrolled release to the environment.

# Target Motion and Cryostat

The Qweak Target base system shall be repurposed for the tritium target. This target system provided reliable motion and cryogenic service for two calendar years. The system also has a base flange that is fully compatible with the Hall A Scattering Chamber. The alterations required to repurpose the Qweak system are summarized below:

- Disconnect and remove all Qweak loop components

- Disconnect/remove all electrical feedthroughs

- Remove support pipe and replace with proper length pipe.

- Install standard alignment mechanism.

- Disconnect cryogenic piping and remove one 4K JT/Bayonet riser section.

- Install new riser with two JT valves.

- Fit/connect refrigerant piping

- Install target mounts and assemblies

Installation of the tritium cell assembly shall be left as a final step prior to closing and evacuating the scattering chamber in preparation for beam (see Section 15.2.1).

## Motion System

The motion of the target is achieved through a two axis stepper motor driven system. The motion axes are perpendicular to the beam in the vertical (y) direction and the horizontal (x) direction with the (z) direction along the beam line. Positioning of the target is accurate to 0.1 mm and repeatable to better than 0.05 mm. The system functioned well for Qweak and is more than adequate in range and accuracy for the tritium target. The two axis positioning allows the target to be accurately centered on the beam line remotely. Thus, motion associated with vacuum and thermal loading can be corrected for. The horizontal (x) system need only be activated during startup (after a cryo-cycle) and shall be deactivated once the beam position on the target has been adjusted for using the alignment target.

Control of the motion system shall be performed via the standard EPICS control system as discussed in in Section 9.

## Cryostat

Details of the altered cryostat can be found in P&ID (JLAB Drawing TGT-103-1001-0000). A schematic view is shown in Figure 4. The system shall use standard components (e.g. JLAB JT valves, bayonets, etc.) and supply 15K He gas from the ESR to cool the target ladder. The cold gas is warmed by the heater to 40K prior to entering the heat sink. The hot return gas is mixed with bypass flow and returned to the ESR cold return with a temperature below 25K. Details of the control system are given in Section 9.

Figure 4: Basic schematic of cryostat.

Figure 5:  Cryostat alarm logic schematic

# Target Cell

## Design

The cell assembly is the primary containment vessel for the tritium gas. It must have a flexible design which safely contains the tritium gas in all modes of operation. These modes are

- General operations in beam

- Shipping to and from JLAB

- Installation and removal at JLAB

- Filling at SRTE.

While installed in the beam at JLAB, the cell must have thin sections which allow the beam in recoil particles to enter and exit the cell with minimal effect from the aluminum walls. While being shipped, these thin sections should be protected by removable covers as should the valve stem assembly. A further challenge to the design, occurs when filling. The normal room temperature fill pressure is 200 psi  [5]. The design pressure while the cell is connected to the filling equipment at SRTE must be above 1000 psi. To meet this later requirement, the covers used for shipping shall also be employed as stays for the thin sections of the cell.  The cell assembly without shipping/filling covers is shown in Figure 6 and Figure 7. The assembly with filling covers is shown in Figure 8. The final shipping assembly where the valve stem assembly is also protected is shown in Figure 9.

Figure 6:  Cell assembly as installed from beam right side. The electron beam is shown as the dashed line and impinges from left to right.

Figure 7:  Cell cross section. Note thin section at tip of main body. The wall thickness exclusive of the beam entrance and exit is 0.018 inch.

Figure 8:  Cell assembly with fill covers installed. Note that the valve is fully accessible for filling/recovery operations.

Figure 9:  Cell assembly with shipping covers installed. Note that the valve is now protected.

### Materials

The cell main body and entrance tube is fabricated from Aluminum 7075-T651 plate (ASTM B209). The fill tube, and flange are fabricated from ASTM A240 SST 304/304L plate. The valve is a Swagelok SS-4BW-BW4-C5-SC11 (modified to have a nut on the actuator) stainless steel 316L ASTM A479. The valve is attached to the fill tube via GTAW butt weld. Swagelok lists the valve working pressure maximum at 1000 psi  [6]. This coupled with a hand calculation shows that the component is under low stress when installed and operated at JLAB at less than 50 psi when cold. We will exercise JLAB policy  [7] to use ASME VIII D1 UHA 51(g) (see Ref  [8]) to exempt the material of the valve from impact testing.

The material used to construct the entrance window and main body is not listed. The following basis was used to determine the allowable stress where ASTM B209 was used to determine the specified minimum tensile and yield.

Sut=72 ksi
Sy=61 ksi
Sa=min13Sut,23Sy=24 ksi

To meet these requirements shipping/filling covers have been designed. Analysis and testing of this final assembly has not been completed as parts are still in fabrication.

### Static Pressure Design

The cell assembly, without shipping/filling covers, has three main components.  It is shown in Figure 6 and Figure 7. For the operational room temperature pressure of 200 psi, the design is adequate.  Formal calculations in compliance with ASME B31.3 304.7.2 (2014) indicate a suitable safety margin (i.e. above 10)  [9].  Burst testing of prototype cells shows a safety factor of nearly 15. These calculations are given in TGT-CALC-103-002.

### Thermal and Cyclic Design

While the above calculations and tests indicate a conservative design, a detailed analysis must include thermal and fatigue effects. In a departure from normal cry target operations, the cell windows affected by the beam see larger temperature fluctuations during beam trips (nearly 85K locally). The pressure in the cell also fluctuates with beam trips but, loads to a lesser extent. During the course of the full experimental run, the tritium cell is expected to experience nearly 18,000 beam trip cycles. The cell is also expected to experience 20 cryogenic cycles from ESR trips and planned shutdowns. A fatigue screening analysis in compliance with VIII Div 2 has also been completed. This calculation is given in TGT-CALC-103-002. This analysis indicates that the depth of the pressure cycle is not of a sufficient magnitude to require a fatigue analysis.

The cell is also thermally loaded. The beam deposits about 3W of heat into the thin aluminum section at the entrance to the cell and similarly 3.3W at the exit. For reasons discussed in Section 6.3, it is required that the cell temperature not exceed 150K while the beam is on. A thermal analysis of the entrance window is given in TGT-CALC-103-008. A similar analysis was performed for the main body in TGT-CALC-103-013.  The results of these calculations are shown in Figure 11 and Figure 11. In summary, the maximum temperature expected on the aluminum wall is about 125K for 20 µA and a raster of 2x2 mm2.

Because the beam generates a local hot spot (and is cyclic), a formal thermal screening was also performed in TGT-CALC-103-002. This analysis shows that the thermal gradient is not large enough warrant a full fatigue analysis. Although not required by Code, a multistep finite element analysis (elastic-plastic) including both thermal and pressure loads was also performed. This analysis is given in TGT-CALC-103-012 and TGT-CALC-130-014. These calculations assumed a pressure load 2 times the room temperature load and full current. The models of the very conservative unphysical loading conditions both indicate a fully acceptable design with maximum equivalent stresses about ½ of the Code allowable for tension. See Figure 12 and Figure 13.

As a further study the effects of the beam at full current when the raster is off are important to understand. Should the effects of such a loading either melt the window or reduce the strength of the aluminum window, a tritium release could occur. An FSD on the raster does exist but, a few milliseconds are required to turn off the beam. A time step analysis of this condition (on the exit window) with full beam current and a beam width of 150 microns (square) is given in TGT-CALC-103-015. The results indicate a maximum temperature on the window of about 206K after 10s which seems to be the maximum (see Figure 14 and Figure 15). While the cell is not adversely affected by a short term raster off event, this condition shall not be allowed.

Note that the material model for the Al 7075 is from the Material Properties Database and includes a multi-linear stress-strain curve and f-N curves at room temperature and -325F.

Figure 10: Thermal analysis of the cell with 20 µA and a 2x2 mm2 raster. Note that the maximum temperature developed on the entrance window is about 120K.

Figure 11:  Thermal analysis of the cell with 20 µA and a 2x2 mm2 raster. Note that the maximum temperature developed on the entrance window is about 120K. The maximum temperature developed on the entire assembly is about 125K and occurs on the beam exit of the cell.

Figure 12:  Equivalent stress on cell entrance window for both pressure loading (400 psi) and thermal loading (beam at full current raster on).

Figure 13:  :  Equivalent stress on cell exit window for both pressure loading (400 psi) and thermal loading (beam at full current raster on).

Figure 14:  Transient thermal model of a raster off event after 10s.

Figure 15:  Temperature of the cell exit window during a raster off event at full current. After 10s the temperature maximum is about 205K.

### Filling and Shipping

Filling and shipping present different challenges.  During shipping and handling, the cell must be safeguarded from impact loading.  The filling process has a system relief set point of 1000 psi necessitating a cell design pressure of at least 1000 psi.  Calculations in compliance with ASME B31.3 (2014) are shown in TGT-CALC-103-002. These calculations indicate that a design pressure of 675 psi is acceptable and does meet Code. The design pressure of 1000 psi required for the filling process cannot be met by the cell alone. To meet these requirements shipping/filling covers have been designed. These covers act as stays on the thin sections of the cells and additionally provide protection from physical impact damage. A formal analysis of the cell and cover assembly is given in TGT-CALC-103-007. The elastic plastic model (see Ref  [8] Section VIII D2) shows that a design pressure of 1000 psi for the cell/cover assembly is acceptable. See

Figure 16: Equivalent stress from elastic-plastic model of the cell with shipping covers installed. The applied load was 3000 psi. See TGT-CALC-103-007.

Hydrostatic testing has been performed on two test assemblies with filling covers installed (the valve port was used to fill pressurize with water). These assemblies were tested until failure by ATS. The seal on Assembly (1) was damaged during assembly by and developed a leak at a pressure slightly above 4000 psi. Assembly (2) failed above 5500 psi. See the figures below.

JLAB shall provide detailed analysis and testing data to SRS/SRNL in all formats required for the SRTE safety basis.

Figure 17:  Assembly 1 following a hydrostatic test. Circled areas indicate damaged sealing surface. Failure pressure (by leak) was 4075 psi.

Figure 18:  Assembly 2 following a hydrostatic test. The cell failed catastrophically at the knuckle region of the hemi head. Failure pressure was 5646 psi.

### Prototype cells

Prototype cells have undergone hydrostatic proof testing to destruction (without shipping covers). The burst pressures for the entrance windows are above 2900 psi. The burst pressures for the main bodies are above 3400 psi. See figures below for details.

Figure 19:  Hydrostatic proof test of entrance window. Window failed slightly above 2900 psi.

Figure 20:  Result of hydrostatic proof test on prototype cell main body. Burst pressure was in excess of 3400 psi.

### Other Gas Cells

The operating room temperature pressure of the cells holding other gasses may be up to 500 psi. This still gives conservatism beyond even Code requirements. It also meets the cyclic screening requirements obviating a detailed fatigue analysis.

## Tritium Permeation

Tritium permeation through the thin sections of the cell wall is expected.  Calculations estimating the amount of permeation through the cell walls are shown in Refs  [10] [5]. The permeation through the cell walls may be determined by  [11]

q=χAlCHDAt

Where

D=D0e-QRT
CH=C0Ppatme-ΔHsRT
A is the surface are
t is cell wall thickness
χAl=molar density of Al
D0 is a pre-exponential factor
C0=1.843 ×10-9
ΔHs=1500 J/mol Heat of reaction

To perform the calculation (TGT-CALC-103-010) the following assumptions were made.

- The solubility of hydrogen (H) in aluminum (Al) alloy 7075 is 10-9.

- The heat of reaction is from Ref  [12] determined from data given therein.

- The diffusion coefficient D is extrapolated from data given in  [13]

- The cell at room temperature for the entire run (this is a conservative assumption)

- The permeation for tritium is 1/3 times that of hydrogen (i.e. scaled by root of molar mass)

Permeation (or leaking) through the Conflat seals and valve stem is equivalent to that of helium.

Over the time period of one year, the estimated tritium loss is expected to be 0.55 Ci  [5,10]. Note that this model is very conservative in that no correction for the lower operating temperature and pressure is accounted for.

## Tritium Embrittlement and Corrosion

It has been shown that the aluminum, at the design temperatures and pressures for the tritium cell is not subjected to corrosion and embrittlement (hydrogen or tritium/He-3)  [14] [15] [16]. However, it has been shown that electron beam impingement on aluminum in a hydrogen environment may induce corrosion  [14] [17].  The authors in Ref  [17] also indicate that the onset of beam induced corrosion is not observed below aluminum temperatures of 180K.  Indeed, this phenomenon has not been observed in any JLAB hydrogen target even gas cell tests performed above 30K. Thus, this is not considered to be an issue as long as the aluminum cell wall temperatures are below 150K when the beam is on.

Tritium embrittlement is also not expected at these pressures and temperatures in aluminum or aluminum alloys  [15].  The presence of the electron beam however increases the fugacity of the tritium in the cell and must be considered. The fugacity is increased by the dissociation of the tritium gas T2→T+T induced by the electron beam. Without the presence of the electron beam, the tritium would dissociate on the wall of the cell as a function of pressure (and other things). Thus, atomic tritium can move from the wall through the lattice by “jumping” between interstitial sites. Increasing the population of atomic tritium can effectively increase the effective pressure (fugacity) of tritium in the cell thereby increasing the concentration of tritium in the lattice. Calculations (TGT-CALC-103-010 and Ref  [18]) have been performed indicating a concentration of tritium in the aluminum as shown in the table below.

| Conditions | Concentration CH |
| --- | --- |
| Room temperature 200 psi in cell | 3.7E-9 |
| Cold with beam on (150K) | 1.5E-8 |

The room temperature pressure required to induce a similar concentration to that of the beam on condition is about 3000 psi. While studies of aluminum at this pressure indicate that is not an issue,  studies are underway between JLAB and SRS/SRNL to examine the short term effects of tritium on this aluminum alloy at elevated pressures and room temperatures  [19].

A further complication is the fact that the tritium in the aluminum will decay over time and trap He-3 atoms in the lattice. A report by Savanah River  [16] indicates that the swelling threshold for He-3 in the lattice of aluminum lithium alloy occurs at a concentration of 4.5E-3 He/Al (at elevated temperature). Assuming, for our conditions, 100% of the tritium in the aluminum converts to He-3 this still allows for a factor 1,000,000 before this threshold is reached.

# Cell Assembly and Installation/Removal

This section describes the cell assembly procedure. Prior to assembly, the individual components which make up the cell shall be both visually and dimensionally inspected as well as cleaned. The welds on the valve stem assembly (JLAB drawing TGT-103-1000-0011) shall be examined and inspected as given on the drawing. After the cell is fully assembled, it shall be leak tested in both normal and reverse modes. Inspection and testing data shall be documented on the cell inspection form. JLAB Target Group Procedures TGT-PROC-15-002 and TGT-PROC-15-003 shall govern the cleaning and assembly of the cell. These procedures are outlined in the following sections.

## Cleaning

The target cell shall be cleaned prior to shipment to SRTE for filling. The components that make up the cell assembly are machined from plate and contaminates from the machine process (e.g. cutting fluids etc.) are expected. These contaminates must be removed. The cell components shall be cleaned using the following procedure:

- Precleaning: Inspect parts for defects and heavy layers of oil and other contaminates. Swagelok components with SC-06 specifications (i.e. oxygen service cleaned) still in original packaging shall be accepted as cleaned. Remove heavy layers of contaminates by wiping with cloth/rag.

- Treat greased/oiled components with Ensolv or equivalent. Repeat this process until surface contaminates cannot be detected visually or by touch.

- Clean all components thoroughly with detergent soap and water. Use bottle brushes, agitators, and abrasives as needed.

- Rinse all parts thoroughly with DI water.

- Place all parts in a previously cleaned stainless steel container suitable for the parts cleaner. Add isopropyl alcohol (99.9%) to a suitable level.

- Place SST container in the ultrasonic cleaner with a DI water bath. Clean parts for 1 hr.

- Rinse parts in DI water.

- Inspect all components thoroughly using VT techniques and JLAB procedure NDT-PQ-100. Use magnifiers and proper lighting.

- Place parts in a dry previously cleaned stainless container.

- Place container in oven at 60 C for one hour.

- Bag individual parts and seal. Store parts in desiccant storage boxes above shop cleanroom.

## Inspection

The components that make up the cell shall be inspected to ensure that thicknesses, materials, surface defects, etc. are within tolerance and specification. These requirements can be found on JLAB drawing TGT-103-1000-0013 and referenced drawings therein. A complete visual inspection of the components shall be performed. Thicknesses of the thin sections of the components shall be made using the MagnaMike system (after it has been calibrated for use). Observations shall be detailed on the inspection form for each cell. The completed form shall be filed in the pressure system folder.

## Assembly

The cell shall be assembled by experienced JLAB Target Group personnel only. Cleaning and inspection shall be performed prior to assembly. The cell assembly procedure is given below.

- Ensure cleaning and inspections have been performed

- Ensure welding examinations and inspections have been performed.

- Install entrance window onto main body.

  - Ensure CF seal is metal to metal. Use A286 12 pt 1” screws and torque all to 140 in-lbf

- Install end cap covers and torque to ¼-28 by 1.5” 12 pt A286 screws to 70 in-lbf

- Install side covers and torque A286 #8-32 x .75 screws to 50 in-lbf

- Install entrance window support with #4-40 A286 x 0.5” screws and torque to 10 in-lbf

- Install the valve support bracket with SST 18-8 #10-32 screws (torque to 50 in-lbf)

- Install the valve stem assembly using A286 8-32 x 0.625” HSHCS torque to 50 in-lbf.

- Examine all mechanical joints visually.

- Record all torque and assembly data on the cell inspection report.

## Testing

The cell assembly shall undergo two leak/pressure tests.

- Leak test with helium leak detector on test vacuum volume with cell assembly inside at pressure. This is known as a reverse mode helium leak test.

- Leak test with cell assembly under pressure as in the above conditions with the valve closed and the VCR cap left off. This test if for determining the valve leak through.

All assembled cells shall be pressure/leak tested with helium in reverse mode using the target group test PIG (the PIG is a test vessel under vacuum with the cell assembly inside). He leaks shall not be observed above 10-9 atm cc/s. A test pressure of 650 psi shall be held for 30 minutes. Testing shall be witnessed by the responsible engineer and performed by a trained technician. Upon completion of this test the cell shall be valved closed, isolating it from the pressure source. The VCR cap shall be left off and the cell assembly retested in the PIG as before. The results of both tests shall be formally recorded on the pressure test form JLAB Form PS-7.

## Cell Purification

The cell shall be cleaned of residual gasses by vacuum pumping using a small turbo pump station and a cold trap. The cell shall be warmed to 100 C using a hot plate. The cell shall be pumped for 48 hours and then back filled with N2 gas.

## Installation and Removal

The tritium cell shall be shipped from SRTE enclosed in an intermediate vessel which is in turn enclosed in a Type A shipping container. Upon receipt of this assembly, it shall be stored in Hall A until it is ready to be installed. Installation of the cell in the scattering chamber shall be the last step prior to closing the chamber for the experimental runs. The following is an outline of the procedural steps required to install the cell:

- Assemble and attach the large access platform to the beam left of the scattering chamber.

- Attach the hut adapter to the chamber.

- Ensure that the motion system set for “tritium in” position and deenergize system.

- Assemble the handling hut and attach to the adapter.

- Connect the scattering chamber large access port located above the beam left thin window to the tritium exhaust system.

- Activate the tritium exhaust system into the flow such that the air velocity at the face of the hut adapter is at least 140 fpm.

- Ensure that Hall A is placed in controlled access.

  - At this point only essential personnel shall be allowed in the hall.

- Move the tritium cell and it shipping containers to the inside of the handling hut.

- Remove the intermediate vessel from the Type A shipping container.

- Remove the Type A shipping container from the handling hut.

- Sample the space inside the intermediate vessel to ensure tritium has not been released from the primary containment, the tritium cell.

- If the sample tritium level is acceptable continue. Otherwise, return the intermediate vessel Type A shipping container and return to SRTE.

- Install the tritium cell in the scattering chamber by following the next steps.

  - Remove the valve in covers and screws from the beam left main body shipping cover.

  - Remove the entrance window plug.

  - Install the small cell mounted collimator.

  - Install alignment pins in the cell.

  - Attach a cell to the copper heat sink in the appropriate location and secure with bolts tightened to 140 in-lbf.

  - Remove the beam left shipping cover and the cover on the exit of the main body.

- Once a cell is installed inside the scattering chamber the hut and adapter may be removed from the side of the chamber.

- The ventilation system shall be kept on until the window cover rollup system is attached.

- Initiate the vacuum system ensure that pump down progressing appropriately.

- The Hall may now be taken into restricted access.

The installation and removal procedures shall be fully developed in the form of an Operational Safety Procedure (OSP).

# Vacuum System

The vacuum system shall provide secondary containment during normal operations and safely handle a full tritium release should the primary containment fail. The scattering chamber, which is the main volume of the system shall also be part of the containment system during installation and removal of the cell from the target ladder (see Section 13). The system consists of the following major components:

- Scattering chamber

  - Main volume of the system (excluding the dump line)

- Main pumping system

  - Two large turbo pumps backed by a mechanical oil sealed pump.

  - Pump exhaust is routed to exhaust system

- Getter system

  - NEG pump with controller and isolation valves

  - NEG pump is backed by a turbo and mechanical pump which are connected to the exhaust system

- Gauges and switches which are connected to the main control system and provide FSD signals

  - Gauges also activate the NEG pump system.

- Beamline isolation window

  - Beryllium window (0.008 in thick) isolates the scattering chamber vacuum from the upstream beamline vacuum space.

  - A small water self-contained water cooling system is used to remove heat from the Be window.

- Dump line

  - The scattering chamber can be isolated from the dump line via a slow acting gate valve.

  - Full time isolation is deemed unnecessary due to existing contamination of this line.

  - The dump line shall be isolated from the chamber during installation, maintenance, and removal operations.

  - Because the dump line is not fully isolated from the scattering chamber, the pumping system for the dump line shall also be routed to the tritium exhaust system.

A schematic representation of the vacuum system is shown in Figure 21. The check valve CK-1 prevents overpressure of the scattering chamber. This valve is also vented to the tritium exhaust system.

Figure 21:  Vacuum system schematic

## Design

The main component of the vacuum system is the scattering chamber which has a volume of about 1900 liters. The standard Hall A scattering chamber (sometimes known as the Big Bite Chamber) is compatible and shall be used. The chamber construction is considered to be acceptable for the pressure and mechanical loads because the Qweak and Hall A Cryogenic target systems are similar in geometry and weight. The chamber shall have two thin sections (windows) on either side of the beam (right and left). These windows have a thickness of 0.014 inch and have been shown to provide an acceptable level of safety for cryogenic hydrogen target systems  [20]. The chamber is shown below in Figure 22.

Figure 22:  Scattering chamber assembly

## Getter System

A getter system (SAES Capacitorr D400-2) with a capacity of 900 Torr-liters of H2 shall be attached to the scattering chamber. The Non-Evaporable Getter (NEG) has a dedicated pumping system (turbo and roughing pumps) that shall be continuously maintaining the getter during normal operations. These pumps shall also be connected to the tritium exhaust system as described in Section 12. The system is shown in Figure 23.

Normal operating conditions are given below:

- Pump valve is open

- Chamber valve is closed

- Pumping system is on

- Getter heater is on

- Ion Gauge is on

Should a vacuum fault be detected on the cold cathode or convectron gauges, the getter system shall activate by switching to the following state:

- Pump valve is closed

- Chamber valve is opened

- Getter heater is off

- Ion gauge is off

- Pump system shall remain on

The system shall be monitored by the main target control system as described in Section 9.

The installation and use of the getter system is considered to be good practice. It shall not be relied on to capture tritium during a release event (see  [21,22]).

Pump valve Norm open

Chamb Valve

Norm Closed

Figure 23: Getter system design. Note roughing pump not shown for clarity. (Tom O’Connor ANL).

## Collimator and Isolation Window

An isolation window is required to separate the upstream beamline and scattering chamber vacuum spaces. The window shall be made from beryllium (99.8%) by Materion Brush Inc. with the following properties.

- IF-1 Brand Be foil EB welded to 2.75” CF flange

- 1.4” disk mounted in 1” aperture

- He leak tight to 10^-9 atm-cc/sec

- Max design temperature 300C

- CF flange is aluminum 2219-T651 (temper prior to welding)

The window shall protrude into the scattering chamber on a reentrant tube which also contains a collimator. The beam will deposit an estimated 3W of heat which shall be removed by water recirculated through copper cooling coils by a NesLab chiller. The coils shall be directly attached to the re-entrant tube. The window shall be positioned roughly 15cm upstream from the entrance window face (thin section). This window is fully capable of supporting a full vacuum load in either direction and will effectively isolate the chamber and upstream beamline vacuum spaces.

The collimator shall be fabricated from HD17, a tungsten copper nickel alloy. It shall slip fit into the tube and have an OD of 1.365 inch an ID of 0.4 inch and a length of 4 inch. This should be more than sufficient to trip the Hall A ion chambers should the beam be steered incorrectly.

The isolation window and collimator assembly is shown in Figure 24.

Figure 24:  Be isolation window and collimator assembly

A thermal model of the Be window in the beam with 30 µA and 2x2 mm2 raster has been given in TGT-CALC-103-009. The maximum expected temperature on the thin window is 32C when the chiller supplies cooling water at 10C. The results are shown in Figure 25.

Figure 25:  Thermal model of the Be window with 30 µA and a 2x2 mm2 raster.

## Alarms

There are several alarms on the vacuum subsystem (see Figure 21). These alarms and their function are listed in the table below. More detailed descriptions can be found in Section 9.

| ALARM | FUNCTION |
| --- | --- |
| PI-1 Cold Cathode Gauge on chamber vacuum | When high issues FSD and activates the getter system And closes main gate valves |
| PAH-1: Pressure switch on chamber vacuum | When high issues FSD and closes main gate valves. Tips audible alarms in Hall and Counting House. |
| FS-1: Flow switch on water cooling circuit for Be window. | When low issues FSD |

Table 2: Vacuum system alarms and functions.

# Control System

## Epics Controls for the JLAB Tritium Target

The hardware and software for controlling the tritium target is based on the scheme that has been successfully used for the Hall A and C cryogenic targets for more than fifteen years.  Specifically, a dedicated target operator in the Hall A counting house will monitor and control the target using a VME-hosted EPICS application.  For the tritium target, EPICS serves the following purposes:

- Monitor various temperature and vacuum levels associated with the target system;

- Maintain a constant target temperature using an automatic, feedback-driven heater;

- Monitor and control the flow of cryogenic helium coolant to the target;

- Control both the vertical and horizontal motion of the target cells;

- Provide a set of alarms to alert users to off-normal target conditions;

- Provide a set of strip charts to track the target performance;

- Archive target performance data;

We do not rely upon the EPICS software to ensure the target system’s safety or integrity.  Instead, hardware and electronic interlocks, described below, are used for these purposes.

The User Interface (UI) is EDM based and offers nearly full control of the target system to the user. Certain aspects of the control system are left to experts however and therefore, have limited access.

## Overview of Target Operation

The Hall C Q-weak target system will be repurposed for the tritium experiments.  This system provides both the necessary vertical and horizontal motions required for the target, and allows easy interface with helium coolant from the End Station Refrigerator (ESR).  An abbreviated P&ID scheme for the target is shown below.

The stack of five gas cells (T2, H2, D2, 3He, Ar) will be cooled using 15 K helium gas from ESR flowing through an copper heat sink attached to the target cells.  Coolant flow through the heat sink will be controlled by Joule-Thompson valve CV-2.  To prevent condensing the gas within these cells, the coolant will be warmed to approximately 40 K by heater H-001 and measured by temperature TS-5.  An automatic feedback (PID) loop between the heater and TS-5 will maintain a consistent coolant temperature +/- 0.1 K.  Thermometer TS-6 on the heat sink will verify its temperature.

For reliable operation of ESR, the helium coolant from the target should be returned to ESR at a temperature not greater than about 25 K.  Therefore, a portion of the flow from ESR will bypass the heater through CV-3 and mix with the 40K flow from the target heat exchanger. A second PID loop, using temperature sensor TS-2, will control the position of CV-3 and ensure the combined coolant is returned to ESR at a sufficiently low temperature.

## Control Hardware and Interlocks

The EPICs application will be hosted on iocha13, the standard VME-based IOC used for the Hall A cryogenic target.  The IOC resides in a VME crate located in the labyrinth outside the experimental hall, while a separate VME crate in the hall contains various IO cards for monitoring and controlling the target.  Repeater cards and an interconnecting cable connect the two crates.  This arrangement has successfully eliminated the radiation-induced failures of the IOC that plagued the early operation of the Hall A cryogenic target.

All system temperatures will be read by a pair of 8-channel, Lake Shore temperature monitors, model LS-218E.  These have been in common use at JLAB for several years.  Communication with the monitors is handled by RS-232 serial connections, and each monitor has a set of user-programmable alarm set points and relays, which will be connected to the FSD system. Redundant temperature sensors and monitors are used to enhance system integrity.  That is, there are actually two thermometers, A and B, installed in each temperature sensor location displayed in the P&I diagram.  The first thermometer of each pair is read by one LS-218 monitor, and the second is read by the other LS-218.  Each monitor will be programmed with identical alarm/relay set points on sensors TS-5 and TS-6, which will trigger a FSD (see Table 1).

The scattering chamber vacuum will be monitored by a cold cathode vacuum gauge, PI-2, and a residual gas analyzer, RGA-1.  The former is a standard cryogenic target device and is read into EPICS using a 16-bit ADC. In case of bad vacuum, the gauge controller for PI-2 will trigger a FSD.  It will also close turbo gate valves CV-5 and CV-6, and open the gate valve CV-7 to getter bed P-006.  While the RGA is not part of the standard cryogenic target instrumentation, the Target Group has successfully used RGAs to monitor the frozen spin polarized target in Hall B.  In that case, a RS-232 connection was made to EPICS, and will be implemented here as well.  We do not foresee any interlocks associated with this device.

The heater power supply will be controlled and monitored by DAC and ADCs, in a manner identical to the Hall A, Hall C, and Qweak cryogenic targets.  A separate power supply will be installed in the Hall A counting house, and can be used to manually control the system temperature in the event of network failure.

The Joule-Thompson (JT) valves and the Warm Return valve to ESR will be actuated and monitored via controllers constructed at JLAB and utilized with cryogenic and polarized targets for many years.  The valves can be controlled manually, either from the experimental hall or the counting house, or via EPICS.  There are no safety interlocks associated with these valves.

The horizontal and vertical motion is achieved by stepper motors controlled by IDC 6961S controllers. The positioning system is controlled though RS-232 serial communication. The accuracy of the system is better than 0.1 mm in each direction and is repeatable to better than 0.05 mm. In practice, the operator of selects individual targets in the UI. Initiating motion issues a FSD and the target moves to a selected location. The FSD is only cleared if the target position read back matches a value in a database accessible only by system experts. The motion system must be calibrated after a cryo-cycle using the beam and the alignment target. This calibration shall be performed only by experts.

We intend to follow the same precautions for moving the tritium target as the Hall A and C cryogenic targets.  The vertical position of the target is determined, with a resolution of about ten microns, by a rotary encoders mounted on the motor shaft and read into EPICS via RS-232.  With the aid of the JLAB Survey and Alignment Group, a unique encoder value for each target cell is determined before the experiment begins.   These values are stored in the EPICS database and are used by MCC to ensure that the target is properly located before beam is sent to the hall.   Horizontal motion is performed using a stepper motor and is likewise monitored by a rotary encoder. Both motor controllers trigger a FSD when motion is first enabled.  Therefore, all target movements must first be announced to MCC, where the FSD can be masked.

The vacuum chamber is isolated from the beam line using a water-cooled, 0.2 mm thick beryllium window.  Water flow switch FS-1 will be installed in the return line from the window and will trigger a FSD if the supply is interrupted.

A summary of the various interlocks and their actions is given in Table 3 below.

| Sensor | Device | Action |
| --- | --- | --- |
| Coolant temperature, TS-4A | Temperature monitor LS-218 #1 | FSD on high alarm |
| Coolant temperature, TS-4B | Temperature monitor LS-218 #2 | FSD on high alarm |
| Heat sink temperature, TS-6A | Temperature monitor LS-218 #1 | FSD on high alarm |
| Heat sink temperature, TS-6B | Temperature monitor LS-218 #2 | FSD on high alarm |
| Scattering chamber vacuum, PI-2 | Cold cathode gauge controller | FSD on high alarm Close CV-5, CV-6 on high alarm Open CV-7 on high alarm |
| Beryllium window cooling water, FS-1 | Water flow switch | FSD on low alarm |
| Tritium detector, T2-ALM-1 | Tritium detector (?) | FSD on high alarm |

Table 3:  Alarms and actions initiated by alarms in control system only

### Machine Fast Shutdown Interlocks

Several multiply redundant interlocks are used to protect the target cells in beryllium isolation window. A schematic of the FSD interlock system is shown in the figure below. Note that these interlocks are all hardware interlocks and do not require software or network to be in operation. In the event of a power failure these interlocks will all fill true and activate the FSD system.

Figure 26:  Schematic of the FSD system for the Tritium Target.

# Shipping

The tritium target cell shall be filled at Savannah River Site (SRTE) must be shipped to Jefferson Lab. The cell shall contain less than 1100 Ci of tritium which may be shipped in a Type A container. The cell shall be shipped inside a secondary container. This container shall be constructed of thick wall stainless steel and be sealed with a 12 inch Conflat flange. This flange shall have sampling ports for detecting leaks from the primary containment vessel (the tritium cell) which may develop as a result of shipping. The secondary container must further be contained in any Type A shipping container which is DOT certified for pressure. This system ensures at least 3 layers of containment during shipping. Upon arrival at Jefferson Lab, the shipping containers shall be only opened in the handling hut immediately prior to installation of the cell. Immediately after opening the outer type a container a tritium detector shall be connected to the sampling port of the intermediate container.

At this time, the requirements for the Type A container which must also meet DOT requirements for compressed gases, have not been fully identified.

# Target Cell Activation

The electron beam will activate the aluminum and alloying elements in the AL 7075 cell.  This activation was calculated using a FLUKA model and summarized in a report  [23].  The expected dose at 30 cm for various decay times is shown in the table below.

| Time after beam operations | Dose mrem/h |
| --- | --- |
| 0 hr | 3 |
| 1 month | 0.3 |
| 1 year | 0.1 |

Table 4:  Expected dose rates at 30 cm from T2 cell after experimental runs.

The main contribution to the expected dose rate is from the following isotopes:  22Na, 65Zn, 51Cr, 7Be, 57Co, 58Co.

Figure 27 :  Expected total dose at 30 cm over time. Plot is from Ref (1).

The figure above indicates that a hold period of 1 month prior to return of the cell to SRS for tritium recovery would be prudent.

# Exhaust System

The exhaust system serves multiple functions. However, the primary function is to remove tritium from Hall A in a controlled manner should the primary containment system fail. The system must function for several different release cases. It must also remove the vacuum system pump exhaust. To minimize the effects of a possible release (see Section 16) the system shall exhaust the contaminated air to a stack. The stack shall have a height 20m above grade at the site boundary. The exhaust system shall serve the following specific functions:

- Exhaust vacuum pumps

  - Exhaust long term tritium from permeation and leakage through seals (<0.5 Ci over run)

  - Remove an acute release should the cell fail.

- Prevent uncontrolled release in the following cases

  - Remove high volume of air from Hall should both primary and secondary containment fail.

  - Supply continuous air removal for handling hut system during installation and removal of the cell. See Section 13.

- Must function as a part of the smoke removal capacity for the Hall in case of fire (required capacity 12000 cfm at 3.5 inH2O)

In order for the exhaust system to function properly, at least one of the truck ramp doors must be closed. Additionally all personnel doors must be in a normally closed state (i.e. not mechanically propped open). Normal personnel access through these doors shall be permitted. A schematic of the exhaust system is shown below. Additional features are shown in JLAB drawing TGT-103-1001-0000 (system P&ID).

## Stack and Blower

The blower shall replace the South East 12000 cfm smoke removal blower located on the Hall A dome. A preliminary design is shown in JLAB drawings A11112-14-01-0100, 0300, 0800 (preliminary). The preliminary design specifies a Penn-Dynamo D30  [24] centrifugal blower with a 15 hp motor. A calculation specifying the air flow requirements and pressure drop is given in TGT-CALC-103-001.  The blower control shall be variable (or two speed) to service the handling hut as well.

The stack must extend roughly 40 ft above the local grade and shall be fabricated from NPS 24 Sch 40 SST 340 pipe. The completed design shall meet the requirements of JLAB FML, Welding and Brazing Program, and ASME STS-1 (2011). The loading factors shall be as given in ASCE-7.

The expected pressure drop for warm moist air at 12000 cfm is 1.74 in H2O, which is the dominant source of pressure drop in the system. With the selected blower however, it is expected that this drop will not have an adverse effect. The design of the system has not yet been finalized and ongoing consulting with SMEs at SRTE shall ensure that the system meets the requirements of the engineering design.

## Vacuum Exhaust

The vacuum exhaust shall be affected through the normal hydrogen target vent system with the exception of the termination point. The current hydrogen exhaust is a NPS 2 Sch 10 SST pipe that terminates at the rear of the Counting House near the Hall A hydrogen storage tanks. The tritium vacuum exhaust will continue this line to the top of the Hall A dome. The vacuum exhaust line shall be strapped to the 24” stack and terminate at the same elevation. This line shall be purged by dry nitrogen at a rate of 1-2 cfm. This is similar to the hydrogen targets which use helium a purge gas.

Figure 28:  Exhaust system Schematic. For more details see the system P&ID drawing sTGT-103-1001-0000.

Figure 29:  Preliminary design of stack/blower system. The 24” stack is fabricated from NPS 24 Sch 40 SST 340 pipe. Note that the stack is supported by guy wires.

## Makeup Air

While the estimated air flow of 525 cfm  [25] required for the handling hut during installation/removal should not present issues to the Hall A makeup air system, the 12000 cfm requirement for tritium exhausting to the atmosphere through the stack is. Such a large air flow (and capacity of the blower) may (provided that there is not sufficient opening for makeup air) could damage the truck ramp door. To ensure that this is not an issue, a louver system on the personnel access door at the bottom of the truck ramp shall be installed. The makeup air can then be supplied through the smoke removal line for the ramp. This would require that the opening at the top of the smoke removal system is not sealed. A calculation TGT-CALC-103-001 shows that the pressure on the overhead door would be acceptable once the louver system is installed. This system shall be part of the personnel door and have dimensions of 2ft x 4 ft. The louver system must be tuned to remain closed until the high speed blower is activated.

# Handling Hut

While the cell is being installed or removed the secondary containment shall be a temporary isolation room fixed to the scattering chamber. This isolation room shall be called the cell handling hut. This hut is being designed and fabricated at SRTE. Access to the hut is provided by a double barrier entrance lock at the rear. The hut shall be directly attached to the scattering chamber via a purpose built adapter with an opening of 15 x 36 inches. The scattering chamber shall be connected to the exhaust system through dedicated ducting. Thus, air shall be pulled through the hut into the chamber and out through the ducting to the stack. An accidental tritium release would then be fully captured inside the hut and chamber and stacked by the exhaust system.

Note:  The hut shall provide secondary containment during installation and removal of the cell.

Air flow rates and pressure drops are discussed in TGT-CALC-103-001. The air flow velocity required to “contain” a tritium release is assumed to be 140 fpm.

## Hut Design

The hut is a temporary structure that shall only be in place during install and removal of the tritium cell. The hut shall be constructed of a clear plastic with access from the rear. The hut shall be “hung” from a frame in a manner similar to other structures designed and fabricated by SRTE. The hut shall be directly attached to the adapter at the front and all seams/joints sealed with tape.

A dedicated platform shall be attached to the pivot deck when the hut is in use. It shall be removed during running conditions. The platform shall have appropriate hand rails and stair access at the rear (side furthest from pivot). The Hall A engineering and technical staffs shall ensure that the structure meets all applicable JLAB design and fabrication requirements.

## Installation and Removal

The hut shall be initially installed by personnel from SRTE instructing those from JLAB. Air flow shall be set and measured at the openings to ensure that the minimum 140 fpm exists at all openings. After the initial test installation, the hut shall be removed and stored in Hall A until needed. The hut shall be stored in a low radiation environment (e.g. upstream of scattering chamber in Hall A) at all times that beam is possible.

# Tritium Detection and Monitoring

A tritium monitoring system shall be developed that includes both fixed and portable monitors. Detection of tritium in Hall A while beam is being delivered, presents a number of challenges. It is difficult to distinguish decays from tritium from the normal background present when the beam is on. There are also no readily available local monitors capable of surviving in the high radiation field present when the beam is on. There shall therefore be two types of devices used for the purpose of tritium detection.

- Local monitors/detectors that are handheld

  - Tyne TAM model 7043

- Remote monitor which samples air pumped from the 1) Hall and 2) stack through a tube to a secure location.

  - Thermo Scientific AMS-4 Beta Air Monitor

The remote monitor must also interface with the FSD and exhaust systems. A schematic of the alarm logic is given in .

Figure 30:  T2 Detector and alarm logic

# Cell Handling

## Filling/Recovery

Both filling and recovery operations shall take place at SRTE.  After the recovery operation is complete the metal cell shall be considered waste. Detailed procedures shall be developed for this operation by SRTE. JLAB shall participate as required in this process.

## Shipping

The cell will be filled with a maximum of 1090 Ci of tritium.  Thus a Type A container may be used for shipping. JLAB is currently working with SRTE shipping and packaging experts to develop a formal shipping method and procedure. We are also working with DOE/DOT. The current plan requires three vessels.

- Tritium cell:  This shall provide the primary containment for tritium while shipping.

- Intermediate Vessel:  This vessel shall provide secondary containment and insulate the cell from mechanical damage that might result from shipping. A custom foam insert shall be inserted to provide this protection. The top flange of this vessel shall also have sampling ports and a protective cover for the valves on the sampling ports.

- Shipping container: This vessel must be approved by the DOT and be suitable for a Type A shipping container. This vessel shall serve as the third level of containment. A portable pressure vessel will meet this requirement although this is excessive.

### Secondary Containment Vessel

The cell shall be suspended from the 12 inch Conflat by mechanical fasteners. The interior of the vessel shall be padded with custom made foam. In this fashion, the cell shall also be heat sunk to the outer Conflat flange. Locking ball valves attached to the 12 inch flange will allow sampling of the atmosphere inside the secondary containment. This will allow a safe assessment of the primary containment, the tritium cell. The sampling ports shall also be capped by Swagelok VCR fittings with anti-tamper devices. The details of the sampling procedure and of the sampling detector are still under development. Should the cell fail inside the vessel the expected pressure rise would be negligible.

### Outer Shipping Vessel

The outer shipping vessel shall be a Type A container. Currently, a DOT pressure vessel has been selected to fulfill this task. This would obviate further approval for shipping a container of pressurized gas.

## Installation and Removal

Handling of the cell outside of the shipping container by JLAB personnel shall be limited to installation and removal only.  These activities shall be performed by fully trained personnel with Hall A in controlled access while all non-essential personnel are prohibited entrance.  The shipping container shall be placed inside a dedicated tritium handling hut with forced air ventilation to the tritium stack.  The cell may only be transferred from or to the shipping container to or from the scattering chamber while the hut is in position and the ventilation (exhaust system) is on (see Section 12).  Installation shall be the last task prior to closing the scattering chamber.  Removal shall be the first task after opening the scattering chamber.  Detailed procedures shall be reviewed and followed. See Section 7.6 for a procedural outline.

# Release of Tritium

There are two possible paths for a release in the tritium target system. It is expected that a small amount of tritium (about 0.5 Ci) will permeate/leak through the thin sections of the cell and the seals on fittings and valve. The second path is from a failure of the primary containment vessel, namely a component(s) on the cell (most likely a thin section). Many steps have been taken to mitigate the risk of full cell failure, however, this path must still be considered as a possibility.

A release of tritium can occur at various phases of operation:

- Case 1:  Shipping

- Case 2:  Installation/removal

- Case 3:Operation

  - A)  Beam permit

  - B)  Controlled Access

  - C)  Restricted Access

- Case 4:  Filling/recovery at SRS

- Case 5:  Packaging and handling at SRS

Each of these cases must be considered separately. The last two cases are considered beyond the scope of this document and shall be considered in the SRTE Safety Basis.

A release of tritium could not only affect JLAB personnel but, the public as well. Several models of release under worst case conditions and a release controlled via the exhaust system and stack have been performed  [21] [22] [26]. Expected dose calculations given in this document are from TGT-CALC-103-004 and TGT-CALC-103-011. The computer code HotSpot  [27] was used to analyze doses to those outside of Hall A.

## Case 1

Given the triple layer of containment and robust nature of the secondary and tertiary containment vessels as well as standard shipping procedures for Type A containers, it is considered extremely unlikely that a release of tritium will occur. Such a release would most likely be caused by a catastrophic failure of the containment vessels from a severe impact load. This failure mode and its effects have been considered in the regulations and procedures regarding the Type A shipment and as such, this Case is given no further consideration at this time.

## Case 2

A release during the installation or removal of the cell would be contained in the handling hut and scattering chamber. A release at this phase could expose personnel to a large dose. The potential causes of such a failure are:

- Dropping cell during handling

- Puncture with tool

- Installation/removal error

To mitigate these causes, procedures have been outlines which ensure that the covers which will protect the thin sections from damage from impact loads remain in place as long as possible. To minimize the dose expected from such an incident the handling hut and exhaust system shall be installed and tested prior to installation. The exhaust system shall also be activated at low speed to ensure that adequate air flow is present (see Section 13). In this fashion, a release would be considered controlled and exhausted through the stack with minimal effect to personnel or the public (see Section 16.4). If a release is considered without the hut or exhaust system on, an exposed worker could reasonably (under very conservative assumptions) receive a maximum dose of 1.2 rem  [21].

If proper procedures are followed for installation and removal, the Hall will be in controlled access with the exhaust system on low power. Cell handling operations must take place in the handling hut. A release under these conditions would then be similar to that of Case 3 A) below.

## Case 3

Each subcase shall be considered separately below.

### Case 3 A)

It is assumed that only the primary containment (cell) fails in beam permit conditions. In this event, the cold cathode vacuum gauge (PI-1) will cause an FSD and activate the getter system. This gauge will trip if the vacuum in the chamber goes above the 10-4 Torr limit which can occur for other reasons (e.g. pump failure, seal loss, etc.). The remote RGA shall be used to determine if a possible T2 release has occurred. The stack monitor shall also be used for this purpose. Under these conditions, a full target release (assuming the getter system is 0% effective) would be exhausted out the stack. See Section 16.4 for details. After the getter has been given sufficient time to collect the tritium, the vacuum system will pumped on for several days to remove residual tritium. After a suitable time, the chamber may be opened for decontamination by wiping down with damp absorbent materials.

### Case 3 B)

A cell failure under these conditions could occur if the cell spontaneously fails while the secondary containment (the scattering chamber) does not; or, more probably, when the chamber vacuum has been breached by a foreign object which also punctures the tritium cell. If procedural steps are properly followed the target shall be in the “HOME” position at all times when personnel are in the Hall (with the exception of installation and removal). This means that the cell would be very difficult to reach with any object such as a pipe. It is conceivable that a serious accident involving a crane or forklift could breach both volumes but, this is considered as extremely unlikely. In such a case, the exposures would be similar to the worst case described in Section 16.2. A release of this type is expected to trigger one or more of the tritium monitors. The events following a monitor trigger are given in Section 14.

In all cases described above, the tritium would be collected by the exhaust system and vented to the atmosphere. Effects from this are given in Section 16.4.

### Case 3 C)

A release while the Hall is in restricted access could generate a worst case exposure scenario for both JLAB personnel and the public should proper procedures not be followed for restricted access. If proper procedures are followed, then the release will have the same effects as those given in Section 16.3.2. Without these procedures, a release through the truck ramp exit is possible. This is the worst case release scenario and is discussed in Section 16.4 as well.

## Effects of Release to Environment Outside of Hall A

An analysis of two release scenarios was performed using HotSpot  [27]. HotSpot is a computer program design for:

- Near surface releases

- Simple weather models

- Short term releases (< 24hr)

- Short range dispersions (< 10 km)

The models are usually conservative in dose predictions. It was assumed that immediately after release, 10% of the tritium converts to HTO while 90% remains HT. This assumption is considered conservative due to measurements summarized in Ref [28] [28].

The models consider both a worst case near ground level release at the exit of the Hal A truck ramp and a controlled release through the exhaust system to the stack. Details of these models are given in TGT-CALC-103-011. A description of the exhaust system is given in Section 12. The results of these models are summarized in Table 5. They show that even in the worst case the dose considered acceptable to the public in the JLAB FSAD is not exceeded. Additionally the worst case dose to JLAB personnel would not exceed 1.2 rem. Models performed by others reach similar conclusions. Results for these calculations are shown in Figure 31 and Figure 32.

| Case | Maximum expected dose (TED) | Distance to max dose | Dose at 300m (TED) |
| --- | --- | --- | --- |
| Truck ramp release | 1.2 rem | 0 m | 4.8 mrem |
| Stack release | 0.96 mrem | 200 m | 0.76 mrem |

Table 5:  Expected maximum dose from various release scenarios and dose at 300 m.

### Assumptions Made in the Models

The following are assumptions made in the HotSpot models

- Release of 1100 Ci 10% is HTO

- Release height 2 m for truck ramp and 20 m for stack.

- 24 hr exposure time

- Exit velocity 0 m/s

- Wind speed at 10 m is 1 m/s

- 500 m inversion layer

- Receptor height 1.5 m

- Stability class F (city)

Figure 31:  HotSpot results for truck ramp release scenario.

Figure 32: HotSpot results for stack release scenario.

# Operational Performance

There is extensive experience at JLAB regarding recirculating cryogenic fluid targets (either gas or liquid). JLAB experience with non-recirculated fluid target is very limited and calibration data was were measured using low beam currents. Administratively, the JLAB limit for aluminum targets (that were not actively cooled by direct fluid contact) is 40 µA. For the tritium target system in general the current limit is 25 µA and for the tritium cell it is lower still at 20 µA. The expected beam heating of the gas in the cell is 2.8W. The linear power density is then 11 mW/mm.  We have seen in other fluid targets that local density (along the beam path) is a strong function of the raster size. A CFD model of the Tritium in the beam at 20 µA with a 2x2 mm2 raster is under development. Material properties relevant to the model have been difficult to obtain. As such these critical quantities (i.e. viscosity, heat capacity, coefficient of thermal conduction, etc.) were derived from a hydrogen real gas model and an ideal gas model. The model assumes that the outer shell is held at 40K with the exception of the beam entrance and exit which are held at 120K. Cooling is fully by convection and is assumed to be laminar. The results of this model as it currently exists are given below. The volume averaged density along the beam is 0.0024 g/cm3. This is about a 20% reduction in density at full current.

Figure 33:  ANSYS CFD model of the tritium space in the beam at 20 µA with a 2x2 mm2 raster. A contour map of the modeled density is shown.

As part of the commissioning of this target, a current scan shall be performed to fully understand the effects of the beam on the target density.

## Tritium Content

In previous discussions, the content of the cell has been assumed to be 100% T2. Tritium has  ½ life of 12.3 years and will decay into He-3. Thus, the pressure in the cell will gradually increase from the moment the fill completes via the decay process. The cell will thus, become immediately contaminated with He-3. This is well known and the data shall be corrected for this.

The initial purity of the tritium content is insured by a redundant mass spectrometer sampling system. The details of this system are not available. A report on the actual tritium purity will be given with the cell at the time of the fill. The purity is expected to be 99.8% pure T2 with contaminates of hydrogen, deuterium and He-3 making up the remainder 0.2%

The precise measurement of the quantity of tritium in the cell is also required. This can be derived from the real gas equation:

PV=nRTZ

Where Z=1.01 is the compressibility of tritium at the fill pressure of 200 psia.

The uncertainties on the quantities above are:

δP=0.2 psia
δV= 0.5 cm3
δT=0.025K

This gives an uncertainty of

δn=1.5%

# Bio-Assay Program

A bio-assay program is considered good practice. All tritium workers and personnel possibly exposed to tritium shall participate in the tritium bio assay program. A tritium worker shall be defined as any personnel involved in the installation or removal of the cell or assembly/disassembly of the scattering chamber rollup cover. JLAB RadCon Group shall develop and administer this program.

# NEPA

An analysis performed by JLAB staff indicates that the potential impacts from the proposed tritium target should not be considered significant within the context of NEPA. The potential impacts associated with the proposed action have been analyzed previously and found not to be significant. The proposed project is determined to be covered under existing NEPA analyses and documentation.  Additional NEPA documentation is therefore not required  [29]. A formal analysis and letter are under development.

# Procedures

Detailed procedure shall be developed for many involving the tritium target system. Procedures for lift plans, general hall work, cryogenic activities etc. are standard and need not be considered here. The following existing procedures shall be modified:

- FSD checkout procedure

  - Modified to include FSDs from tritium target which are additional to the standard cryogenic target FSDs. This procedure shall be controlled by Accelerator protocols.

- Hall checklist

  - A number of modifications shall be made in the procedure regarding the tritium target and preparation for beam. This procedure shall be filed with document control.

  - The remote tritium monitor should be tested somehow prior to lock up. This shall be a RadCon procedure.

- Hut installation and removal

  - This procedure is unique to the tritium target and has been outlined in Section 13.2. It shall be filed in Document Control.

- Tritium target installations and removal

  - This procedure has been outlined in Section 7.6. It shall be filed in Document Control.

- Tritium cell inspection and assembly procedure.

  - Procedure is complete and shall be filed in Document Control

- Component cleaning procedure

  - This procedure is complete and shall be filed in Document Control.

- T2 Exhaust testing procedure

  - This procedure is required to ensure that the exhaust system is functioning correctly in all modes. It shall be filed in Document Control.

- Access to Hall A

  - Restricted access to Hall A shall be limited to personnel with specific tritium training. An outline of the procedure is given below.

  - Controlled access to Hall shall be similar to the current Controlled Access protocols. A procedural outline is given below.

  - These procedures shall be filed as part of the Experimental Physics Documents (e.g. COO, ESAD, RSAD, etc.).

- Target operating procedures shall be developed. These procedures will be similar to the existing cryogenic target procedures. This procedure shall be filed in Document Control.

- T2 alarm checkout procedures

  - JLAB RadCon shall be responsible for these procedures and how they are maintained.

## Hall Access Procedures

Access to the Hall will have to be more strictly controlled while the tritium target is located in Hall A. The Hall itself shall be used as part of the three level containment system and its integrity shall be maintained at all times. This does not mean that all access will be through air locks etc. Regular personnel access through man doors shall be allowed as normal. These doors shall not be propped open. Access to the Hall is provided by the following portals:

- Truck ramp door system which consists of a door located at the bottom of the truck ramp and one at the top. These are large overhead doors that are electronically controlled. Activation of the lower door shall be allowed through the CANS system only. Should maintenance be required on this system, a procedure shall be developed and approved by RadCon and the Tritium Target DA. Opening this door shall activate the high speed exhaust system which shall remain on until this door is closed. Signage shall be posted near the controls for the door outlining the protocols for access.

- Personnel access door at bottom of the ramp. This door shall also be on the CANS system. Louvers shall be installed in this door to provide makeup air for the exhaust system. This door shall not be propped open. There shall be signage posted on this door indicating the training required and the protocols for access.

- Personnel access door from accelerator: Access through this door shall be allowed only in the event of an emergency. It shall remain locked by electronic system.

- Personnel access door from labyrinth: Access through the labyrinth shall be controlled via the CANS system at the first access door. Signage shall be posted on this door describing the protocols for access and the training required for entry. This door shall also be controlled by the MCC/SSO as required during a controlled access. This door and indeed none of the labyrinth doors shall be propped open.

### Truck Ramp Protocols

The truck ramp presents the largest challenge to ensuring that the Hall maintains the third level of containment. The truck ramp door at the bottom of the ramp shall be activated by the CANS system only. The access list for this door shall be very limited. All personnel on this list shall be trained for this specific task. The procedure for truck ramp access shall only be allowed in Restricted Access Mode.

- At the start of this procedure all ramp doors shall be closed.

- For loads entering the Hall:

  - The top ramp door shall be opened by personnel located on the inside of the ramp.

  - The vehicle carrying this load shall enter the truck ramp.

  - The upper door shall be closed and the load carried to the bottom of the ramp.

  - The lower door shall be opened.

    - This activates the high speed exhaust system which should provide sufficient noise to ensure a strong desire to reclose the lower door.

  - If the lower door can be closed (i.e. the truck is not stuck half way) then it shall.

  - The load shall be brought into the Hall and unloaded promptly and safely.

  - Once the load is removed the truck shall exit the Hall Promptly and the lower door closed.

  - The truck may proceed to the top of the ramp where the upper door is opened temporarily to allow its passage.

- For loads leaving the Hall the truck shall enter as above. Upon reaching the interior of the Hall it shall be loaded promptly and exit the Hall forthwith using the same procedure as above.

### Controlled Access

Controlled access to the Hall while the tritium target is located in Hall A shall be similar to the standard controlled access. The following is an outline of additional steps which shall be addressed in the Controlled Access Procedure:

- Make entry as required with the additional training requirements for tritium.

- Ensure that ARM/RadCon has handheld TAM (tritium air monitor). The maintenance of these monitors is the purview of the RadCon group.

- All large truck ramp doors shall remain closed during this access

- Access is allowed to the Hall if an acceptable level of tritium is detected. Note that it is assumed that there will be some airborne tritium from normal beam operations present in the Hall.

- If unacceptable levels of tritium are found then exit immediately and inform RadCon. Activate the exhaust system and await further instructions. All personnel on the access team shall be enrolled in the bio-assay program.

- The following activities are not allowed during a controlled access for these accesses:

  - Forklift operations

  - Lift operations within 20 ft of the pivot

  - Crane operations

  - Heavy work (other than diagnostic and tuning) within 20 ft of the pivot.

- Should the TAM alarm at any point during the access, all personnel shall promptly leave the Hall and activate the exhaust system. Notify RadCon for further requirements.

- Should power fail, all personnel shall promptly leave the Hall.

### Restricted Access

Access shall be allowed under restricted access as normal with the following exceptions:

- The access to the Hall shall be controlled through the CANS system.

- Truck ramp protocols as described above shall be followed.

- To change to or from controlled access the scattering chamber window covers shall be installed.

- All work within 20 ft of the pivot shall be “covered” by a handheld TAM monitor.

- At least on TAM shall be available at the entrance of the Hall for access.

- This must remain located near the pivot. It shall be maintained by the RadCon group.

- Should power fail, all personnel shall promptly leave the Hall.

- Heavy work (crane operations, operations requiring hand tool work directly on the chamber, fork lift operations etc.) performed within 20 ft of the chamber shall be governed by an RWP.

### Special Access for Installation and Removal of Cell

The installation and removal procedure for the tritium cell has been outlined in Section 7.6. The Hall shall be placed in a special access mode similar to controlled access for these procedures. This access mode shall include:

- All pre-installation tasks shall be performed including the placement of tools and materials in preparation for cell installation or removal. A special tritium target tool box shall be maintained for this specific purpose and no other, but the JLAB Target Group.

- The exhaust system shall be activated at low speed and the shipping container holding the cell shall be placed in the handling hut.

- The Hall shall be placed in controlled access. This includes a full sweep and key access.

- Entrance to the Hall shall only be allowed for personnel directly involved in the installation/removal process.

- This activities shall be 100% covered by RadCon operating under a specific RWP.

- Upon completion of the cell installation, the team closing the chamber may enter the hall to perform this specific task while being supported by RadCon and the Target Group.

- The JLAB DA shall be responsible for the cell installation removal.

- Once the chamber is sealed and vacuum is established the Hall may be returned to restricted access.

# Hazard Analysis

The tritium target is an alteration of the existing Qweak Target. The alterations required for this system to function for the tritium target are associated with the motion system (low voltage stepper motors) and cryogenic piping. A number of additional hazards associated with tritium are assumed. A loss of 1100 Ci of tritium are not in fact expected to pose a serious risk to JLAB personnel, the public, or the environment under any reasonable failure scenario. However, it is assumed that any loss of tritium could expose JLAB personnel to doses above those allowed by lab policy and could impose significant decontamination costs and interrupt schedules.

A formal Failure Modes and Effects Analysis using MIL-SPEC-1629A as guidance is given in Ref  [30].

The following hazards have been identified as applicable to the tritium target:

- Pressure hazards

  - Vacuum systems (scattering chamber)

  - Cryogenic piping

  - Tritium gas cell (200 psia)

  - Other gas cells (500 psia)

- Cryogenic hazards

  - Exposure to cold fluids

  - ODH

- Radiological hazards

  - Excessive dose

  - Contamination

  - Activation of materials

The Hall A Cryogenic Target controls (with minor alterations) and scattering chamber are also being used as part of the system. Note that the overpressure protection for the scattering chamber shall remain installed as shall be vented to the tritium exhaust system (on the pump line). A hazard analysis for each of these systems has been performed and is documented in the Hall A Standard Equipment ESAD and Ref  [20]. No further consideration for these systems is given here. An ODH assessment was performed for the cryogenic target and magnets in Hall A. No appreciable hazard was found below the crane level. The ESR supply system for the tritium target system is identical to that of the Hall A cryogenic target. Therefore, no further analysis is performed here.

The following hazards are considered to have been addressed

- Cryogenic hazards (equivalent to Hall A cryogenic target)

  - Exposure to cold fluids

  - ODH

- Vacuum hazards (Equivalent to Hall A cryogenic target)

- Electrical hazards (equivalent (or lower) than Hall A target

- Material handling hazards (equivalent to Hall A target

- Chemical hazards equivalent to Hall A target

- Activation of materials equivalent to Hall A target

JLAB ESH Manual Chapter 3210 Appendix T3 describes the method required for hazard analysis. This method may be clearly applied to individual tasks but is difficult to apply to a system. The standard ANSI/GEIA-STD-0010  [32] shall be used to determine the probability of a given incident.

| Likelihood | JLAB Level | Number assigned in FMEA | Definition |
| --- | --- | --- | --- |
| Probable | H | 4 | Probability of occurrence per operational hour is greater than 1E-5 |
| Remote | M | 3 | Probability of occurrence per operational hour is less than 1E-5, but greater than 1E-7 |
| Extremely Remote | L | 2 | Probability of occurrence per operational hour is less than 1E-7 but greater than 1E-9 |
| Extremely Improbable | EL | 1 | Probability of occurrence per operational hour is less than 1E-9 |

Table 6:  ANSI/GEIA-STD-0010 Standard likelihood of occurrence  [32].

## Pressure hazards

The pressure hazards applicable to the tritium target are given below in Table 7.

| Hazard | Unmitigated Risk Code | Mitigation | Mitigated Risk Code |
| --- | --- | --- | --- |
| Cell failure Flying debris Loud Noise/pressure wave | 2 | -Designed to Code (675 psi design pressure) -Procedures/training -PPE -Shipping covers on until final step of installation -Low stored energy in pressure from small volume. | 1 |
| Cryogenic piping: Explosion flying debris Exposure to cold gas ODH Load noise/pressure wave | 4 | -Design/fabricated/examined/tested to Code. -Proper/adequate relief is installed. -Procedures/training | 1 |

Table 7: Pressure hazards with mitigations

## Radiological Hazards

Tritium presents a number of atypical hazards to JLAB personnel. The unmitigated risk code associated with tritium shall be 4. This is a conservative assumption based on the potential for public exposure and elevated dose (above JLAB limits) for workers on site. Furthermore, the costs associated with decontamination may be significant if no mitigating steps are taken after a release. The table below addresses the radiological hazards associated with the target system.

| Hazard | Unmitigated Risk Code | Mitigation | Mitigated Risk Code |
| --- | --- | --- | --- |
| Uncontrolled release of 1100 Ci of T2: 1) Ground level release to environment 2) Worker exposure 3) public exposure 4) contamination of Surfaces in Hall A Contamination of beam line 5) hi dose to installation team | 4 | - Design/construct the cell to Code with conservative FS - FSD protections - collimators - Be isolation window - fixed and portable T2 monitors - Interlock system - exhaust system with 20m stack - Procedures/training - simple installation - 3 levels of containment - Getter system | 2 |
| Uncontrolled leak of T2 from cell seals and thin sections: 1) Possible contamination of beamline 2) possible exposure of workers in Hall. 3) low level release of T2 to hall and contamination thereof. | 3 | - All above mitigations - Seals are typically He leak tight even when cold. - Pumps vented to exhaust system | 1 |

# Response to Committee (June 3, 2010) Report

The committee formed on June 3 2010, issued a report from which 45 action items were derived. A response for each item Is given in the following sections.

## Administrative Action Items

### Task 1:

Appoint a lead person at JLAB, a design authority, who will emphasize the engineering aspects of the target system and who will be responsible for the design, fabrication, procurement, installation and operation of the system.

This has been completed. The JLAB Design Authority Is David Meekins.

### Task 2:

Establish an engineering team.

This task has been completed the team members are

- Dave Meekins (Responsible Engineer)

- Hall A Engineering Group (R. Wines lead)

- R. Holt and T. O’Conner ANL

- JLAB Target Group (C. Keith lead)

- SRTE (J. Novajosky manager)

- SRNL (H. Lee Nigg lead)

### Task 3:

Develop clear responsibilities for INL, JLAB, and collaboration; and determine who will sign off on safety checkout plans.

This requirement is no longer directly applicable. There are two current agreements (ICO) with SRS/SRNL to perform the cell filling and material testing  [19] [33]. These agreements clearly define tasks and budgets. SRS personnel shall be responsible for completing the required Safety Basis. JLAB personnel shall provide all necessary calculations and proof testing data to SRS. Additionally JLAB shall provide leak test, weld and material documentation, and inspection assembly reports to SRS with the active cell in preparation for filling.

### Task 4:

Administrative requirement that beam blow-up optics are used in experiment.

A thermal analysis of the cell when the raster is off in full beam current assumes a square beam profile of 0.150 mm width. This model is shown in Section 6.1.3 and shows that the cell is sufficiently designed to survive this condition for a limited time period. Concerns over beam induced hydrogen corrosion indicate that the cell operating temperature must be kept below 180K. The raster off condition does not meet this requirement thus, it is only allowed for very short periods of time prior to the FSD trip on raster.

### Task 5:

Administrative limit on beam trip rate when cell is in the beam.

A full fatigue screening analysis was performed in accordance with ASME BPVC VIII D2 Part 5. This analysis indicates that beam trips occurring at a rate of 15 per hour. A cell duty factor of 30% was assumed (there are 5 other targets). This calculation is given in TGT-CALC-103-002 and indicates that this trip rate does not warrant a formal fatigue analysis due to the low stress of the cycle.

A safety algorithm was developed by R. Holt that also addresses this issue  [34].

### Task 6:

Make a provision that if the H or He target fails, a failure mode determination is made before the experiment continues with tritium gas.

A failure of any other cell will cause the vacuum monitor alarm to trip activating the getter system and many other controls. An RGA shall be installed to give some indication of the species of the gas after such a failure. Vacuum will have to recover prior to using this device. If it is determined that the one of the other cells failed for an undetermined reason the experiments shall be stopped and the tritium cell secured and shipped back to SRTE for recovery. If the failure mode is from an operational or procedural source then the situation can be assessed at the time of failure. A determination to continue the run shall be made by JLAB management.

### Task 7:

Extensive review of final cell design and test results should be performed by JLAB, Argonne, and outside experts.

This process is ongoing. In addition to the institutions listed above SMEs at SRTE/SRNL shall also review the design.

## Target Cell

### Task 1:

Develop a modular target design.

The current design is modular and uses techniques and components similar in design to existing JLAB cryogenic hydrogen and helium targets.

### Task 2:

Use a 1000 Ci source.

The current design specifies a fill of 1099 Ci of tritium. This is under the limit required for Type A shipments. Release and mechanical calculations indicate that this is a safe design. See TGT-CALC-103-002 and TGT-CALC-103-0011.

### Task 3:

Determine maximum target window thickness based on physics requirements.

The operation of the polarized 3He gas target cell at JLAB provides a baseline for this study since this target has many similarities to the present tritium gas target design in terms of length and gas pressure.  It appears that we can use up to three times thicker windows than that of the polarized 3He target without seriously compromising the experiment.  From a Monte Carlo simulation, it appears that there would be less than a 20% loss in target thickness that would be necessary to cut out the target windows from the data.  Thus, the present tritium gas target design uses windows that are approximately a factor of three thicker than those of the polarized 3He target.

### Task 4:

Target cell should comply with ASME Boiler and Pressure Vessel Code, Section VIII Div 2 2007

This is not truly possible at the cell cannot be meet all the requirements of this Code nor can it be stamped. Furthermore, this Code is not the most applicable ASME Code. Using JLAB policy for selecting the proper ASME Pressure Code results in the code selection of ASME B31.3 (2014). This shall be the Code of record. This Code may be followed in full with the exception of impact testing. The JLAB policy for fracture toughness testing shall be followed  [7]. Section 304.7.2 of B31.3 allows for the use of ASME BPVC VIII D2 FEA analysis. This shall be performed using the allowable stress in tension from B31.3.

### Task 5:

Cell should be filled offsite and be designed to survive transport.

The cell has been developed to safely operate with the design pressure of 675 psi. This design pressure includes safety margins required by ASME B31.3. The actual maximum operating pressure for the tritium cell is 200 psi. Therefore an additional factor of safety exists above Code requirements. The cell shall also be fitted with custom (precision fitted to each cell) shipping/filling covers. These covers act as stays for the thin sections (which is allowed by Code). An analysis of the cell assembly with these covers (compliant with ASME B31.3 using FEA and elastic-plastic model TGT-CALC-103-007) shows that the design pressure of 1000 psi is acceptable. These covers also serve to protect the thin sections from impact loading. A proof test was performed on two assemblies using a hydrotest incompliance with ASME BPVC VIII D1 UG-101. The correctly assembled cell burst above 5500 psi. The incorrectly assembled cell did not fail but developed a leak on the damages sealing surface above 4000 psi. We conclude that this design is sufficient for protecting the cell while shipping/handling and while filling.

### Task 6:

Cell should survive cyclic loading (beam trips).

See Section 22.1.5. The cell design is adequate for beam trips.

### Task 7:

The tungsten collimator should be better supported.

The current design of the target system specifies two collimators: 1) a short collimator installed on the face of the cell (that moves with the cell) and 2) a long collimator upstream of the target located in the reentrant tube. This collimator is now fully supported by the tube.

### Task 8:

Valves should have all metal wetted parts.

Selected valves are Swagelok all metal bellows sealed valve. Specification is also in accordance with SRTE requirements.

### Task 9:

Cell must sustain a full vacuum load.

The cell is of sufficient design to with stand a full vacuum load.

### Task 10:

Target system should be designed to be cooled with 90K N2.

The target system has been designed to use the ESR supplied 15K coolant which shall be heated to 40K. This will provide more than sufficient coolant to prevent the cell walls from reaching 180K which is the temperature of concern. See Section 22.1.4.

### Task 11:

Verify that Al 2219 is a suitable material for tritium gas.

The current cell material is Al 7075. This aluminum alloy is shall be considered acceptable for tritium use. See Section .

### Task 12:

Determine strength of welded Al 2219.

This is obviated by the current design.

### Task 13:

Weld coupons should be tensile and bend tested

See above.

### Task 14:

Al –ss transition piece should be purchased and elbows should be used

This is obviated by the current design.

### Task 15:

An elastic plastic model of the cell (ASME D2 5.2.4) should be used.

Completed in the following calculations: TGT-CALC-103-002 -007 -012 and -014. These models indicate that the design is adequate for all conditions.

### Task 16:

Proof tests on more than 3 endcaps and at least one complete cell should be performed.

Testing has been performed on multiple assemblies including assemblies with shipping covers installed. These tests all indicate a conservative design.

### Task 17:

Heat cycling tests with a tritium loaded cell should be performed for a period of 6 months.

This action item creates challenges that are difficult to overcome and of results of questionable value. Note that to pressure test the exposed cell would require removing the tritium that it contains. Any embrittlement that would be present (many references indicate that there would be none See Section 6.3) would be reversed by the removal the gas. The exception to this reversible embrittlement is swelling (or blistering activated by He3 in the lattice). Studies at SRS indicate that we are many orders of magnitude below the concentration required for this process to initiate. Nonetheless, a study is underway at SRTE/SRNL to investigate the fatigue crack growth of Aluminum 7075 at pressures exceeding 3000 psi. This is 15 times higher than the room temperature design pressure.  Data from this study shall be collected every 4 months for one year  [33]. See Section 6.3 for more detail.

### Task 18:

Consult Wayne Kanady at INL and Walter Shmayda at University of Rochester regarding tritium diffusion.

Several models for tritium diffusion were developed  [5,10]. The estimated diffusion for the cell at room temperature over 1 year is 0.5 Ci. This is conservative estimate because the cell will be cold and at much lower pressure much of this time.   R. Ricker at NIST and Brian Somerday at Sandia were consulted and this resulted in a report on the diffusion and permeability in the target cell.

### Task 19:

If insufficient information exists on beam shock wave tests, perform such tests at JLAB.

Cyclic loading from beam trips was determined not to be an issue  [9,35,36].

### Task 20

Cell must survive off-normal beam conditions for at least 3 times the amount of time that it takes for an interlock to turn the beam off.

The cell can survive more than 10 seconds with no raster. See Section 6.1.3.

### Task 1

Determine DOT and DOE regulations for shipping filled target cells to JLAB.

This process is ongoing at this time. JLAB is consulting SRTE shipping and packaging SMEs as well as DOE/DOT for shipping requirements for small quantities of compressed gas. Three layers of containment shall be provided in the final solution.

### Task 2

Secondary containment should be physically isolated from beamline.

The beam line is isolated by a beryllium window. See Section 8.3.

### Task 3

The scattering chamber pumps shall be vented through tritium stack.

All pumps including the scattering chamber, getter, and dump line pumps shall be exhausted through the tritium exhaust system to the stack. See Section 8.

### Task 4

The scattering chamber should be monitored for high and low levels of tritium.

True real time monitoring of the tritium levels in the chamber may be unrealistic given the high radiation field in the Hall when the beam is on. Multiple vacuum gauges shall be employed however. If the vacuum level rises the first response will be to assume that this rise is due to tritium release. An RGA system shall be connected to the chamber as well which should aid in the diagnosis of the vacuum rise. The stack monitor will also indicate tritium if the vacuum pumps are removing tritium from the chamber. The source of the vacuum failure shall be assessed before continuing with repair procedures.

### Task 5

A U getter bed should be attached to the scattering chamber.

To avoid hazard associated with uranium, a getter from SAES systems shall be used. This getter is described in Section 8.2.

### Task 6

An additional long collimator should be placed upstream.

A collimator shall be placed on each of the target cells (that moves with the target) and a second long collimator shall be placed just upstream of the Be isolation window.

### Task 7

Dedicated vent pumps/fans and lines should be installed over the scattering chamber and used for installation and removal procedures.

See Section 12.

### Task 8

Airborne radioactivity detectors interlocked with the vent/fan stack system should be used.

The system design incorporates this. See Section 12.

### Task 9

Manual scram buttons in the hall and counting house for the ventilation.

### Task 10

Additional beam raster detector and interlock system.

The fast raster now has a FSD on a current comparator in addition to the FSD on the power supply enable signal. This system is currently in use.

## EHS Tasks

### Task 1

Establish baseline for detectable tritium at the JLAB site.

See Ref  [37,38] for basis and details.

1. Tritium in water
- We can readily detect tritium in water at about 500 pCi/L.  The EPA drinking water limit is 20,000 pCi/L.  Environmental permits for groundwater and stormwater effluent require a minimum detectable activity (MDA) of 1000 pCi/L.

2. Tritium as surface contamination
- The DOE surface contamination limit for tritium is 10,000 dpm/100 cm^2.  We can easily detect surface contamination at levels 100 times below this.

3. The manufacturer reports that the air monitor JLAB has will see 10^-7 uCi/cc of tritium.  The DAC for elemental tritium is 2E-1 uCi/cc and for HTO is 2E-5 uCi/cc.  So, even if all the activity was HTO, the monitor can see it at less than 1% of the DAC.

4. According to the spec sheets, the portable samplers that SRS uses (Sartrex 209L) will see 1E-5 uCi/cc tritium.  We expect to use these during transfer of the target.

5. Grab air samples can also be taken for HTO using bubblers. This is not expected to be used as a primary sampling technique, due to the T2 source material but, we may employ this method as a confirmatory process.  MDA for this technique is usually in the 10^-8 uCi/cc range or lower.
 
6. We are currently working on analysis techniques to measure tritium in oils and other material such as concrete.  These are not directly related to the operational aspects of the tritium target, but are part of our efforts to increase our ability to make measurements to confirm absence of radioactivity in materials released from control.  The capability to measure tritium in such materials should not be thought of as a necessary component of the program, as I doubt that even SRS can do it (they probably send these kinds of materials out for analysis, if needed).

7. Based on feedback from SRS, and our own procedures and ALARA concerns, we will run some sort of H3 spot check bioassay.  We would probably get these analyzed offsite (mainly because of the biohazard issues handling urine).  Typical sensitivity for urine is like our sensitivity in water (less than 1000 pCi/L).  At 1000 pCi/L in the urine, the dose is about 2E-4 mrem per day.

### Task 2

Develop algorithm for safety involving amount of tritium, beam current, beam time.

Complete. See Tritium Technical Report Ref  [34].

### Task 3

Worst-case scenarios for worker exposure and all dose calculations should be analyzed or calculated by qualified personnel.

Several models for tritium release have been developed by qualified personnel. See Section 16.

### Task 4

A more detailed assessment of impact of target loss on Hall A should be performed by qualified personnel.

Several models for tritium release have been developed by qualified personnel. See Section 16.

### Task 5

A more detailed assessment of impact of target loss on Hall A should be performed by qualified personnel.

Several models for tritium release have been developed by qualified personnel. See Section 16.

### Task 6

Use the ICRP-68 dose coefficient of 1.8E-11 Sv/Bq reference for exposure evaluations.

The requirement that the released tritium is immediately converted to HTO is overly conservative. Many measurements  [39] indicate that the conversion is typically less than 1% per day. The rate assumed for conversion is 10% in current calculations. Note that several models for tritium release have been developed by qualified personnel. See Section 16.

### Task 7

Use 10 mrem as the maximum allowed site boundary dose.

Several models for tritium release have been developed by qualified personnel. See Section 16. The expected dose at the site boundary is less than 1 mrem for a stacked release. For the worst case scenario, the expected dose at the nearest site boundary is slightly above the limit (with the conservative 10% HTO assumption). The expected dose at 300 meters is however less than the 10 mrem limit.

### Task 8

The risk analysis should follow tables 4.2-4.5 of JLAB’s FSAD, rev. 6 and use realistic target failure probabilities.

See Ref  [40] and Section 20.

# Summary of Calculations

Below is a list of calculations supporting some of the data used in this document.

Table 23-1: List of Completed Calculations

The following is an embedded link to the calculations listed above.

# References

[1]	J. Arrington, D. F. Geesaman, K. Hafidi, R. J. Holt, D. Potterveld, P. Reimer, J. Rubin, J. Singh, X. Zhan, K. A. Aniol, D. J. Margaziotis, M. B. Epstein, G. Fanourakis, J. Annand, D. Ireland, R. Kaiser, G. Rosner, E. Cisbani, F. Cussano, S. Frullani, F. Garibaldi, M. Iodice, L. Lagamba, R. De Leo, E. Pace, G. M. Urciuoli, E. Chudakov, J. Gomez, D. W. Higinbotham, C. W. De Jager, J. Lerose, D. Meekins, W. Melnitchouk, R. Michaels, S. K. Nanda, B. Sawatsky, P. Solvignon, A. Saha, B. Wojtsekhowski, B. D. Anderson, A. T. Katramatou, D. M. Manley, S. Margetis, G. G. Petratos, W. Korsch, X. Jiang, A. Puckett, E. Beise, J. R. Calarco, K. Slifer, and S. Scopetta, 1 (2010).

[2]	C. Hyde, S. E. Kuhn, L. B. Weinstein, I. Korover, J. Lichtenstadt, E. Piasetzky, I. Yaron, T. Aviv, and J. W. Watson, PR12-14-011 Proton and Neutron Momentum Distributions in A = 3 Asymmetric Nuclei A Hall A Collaboration Experiment Proposal PR12-13-012 to Jefferson Lab PAC 42 , July 2014 (2014).

[3]	L. S. Myers, D. W. Higinbotham, and J. R. Arrington, Ratio of the Electric Form Factor in the Mirror Nuclei He-3 and H-3 (2014).

[4]	P. Solvignon and E. Al., Precision Measurement of the Isospin Dependence in the 2N and 3N Short Range Correlation Region (n.d.).

[5]	D. Meekins, Estimated Tritium Pressure, Solubility, and Permeation in Cell (2015).

[6]	Swagelok, 1 (n.d.).

[7]	D. Meekins and et. al., Policy for Fracture Toughness Testing Requirements for Pressure Systems and Components at Low Cryogenic Temperatures (2010).

[8]	ASME, ASME Boiler and Pressure Vessel Code. (2013).

[9]	D. Meekins, Tritium Target General Target Cell Calculations (Code) (2015).

[10]	R. J. Holt, R. E. Ricker, and D. Meekins, Tritium Permeability of the Al Target Cell (2012).

[11]	E. Cussler, Diffusion: Mass Transfer in Fluid Systems (2009).

[12]	J. R. Scully, G. a. Young, and S. W. Smith, Mater. Sci. Forums 331-337, 1583 (2008).

[13]	R. Braun, H. Schluter, H. Zuchner, H. Buhl, and A. Turnbull, Inst. Mater., Teddingt. (1994).

[14]	R. Lasser, (1989).

[15]	A. Troiano, R. Gibala, and R. Hehemann, Hydrogen Embrittlement and Stress Corrosion Cracking: A Troiano Festschrift (1984).

[16]	M. R. Mac Louthan, 31 (2000).

[17]	H. M. Flower, Radiat. Eff. 33, 173 (1977).

[18]	R. J. Holt, R. E. Ricker, D. Meekins, and B. Somerday, Beam-Induced and Tritium-Assisted Embrittlement of the Target Cell at Jefferson Lab (2013).

[19]	M. J. Morgan, H. L. Nigg, and J. D. Novajosky, ( U ) FY15 / 16 TJNAF-SRTE Tritium Target Structure Evaluation (2016).

[20]	M. Seely and D. Meekins, Hydrogen Target Safety Assessment Document (2004).

[21]	D. Meekins, Tritium Target General Release Calculations (2015).

[22]	D. Meekins, HotSpot Analysis of Tritium Release (2015).

[23]	G. Kharashvili, Activation of Tritium Target Cell (2014).

[24]	PennBarry, Model : D Class I Centrifugal Fans (n.d.).

[25]	D. Meekins, Tritium Stack General Calculations for Air Flow (2015).

[26]	B. Napier and R. J. Holt, Analysis of a Tritium Target Release at Jefferson Lab (2012).

[27]	S. Homan, (2015).

[28]	D. G. Jacobs, ORNL (n.d.).

[29]	G. Roemer, JLAB MEMO 1 (2015).

[30]	D. Meekins, Hall A Tritium Target FMEA (2015).

[31]	D. Stamatis, Failure Mode and Effect Analysis: FMEA from Theory to Execution (2003).

[32]	ANSI, ANSI/GEIA-STD-0010 (2010).

[33]	T. S. Mcgee, H. L. Nigg, and J. D. Novajosky, ICO FY15 / 16 TJNAF-SRTE Tritium Experiment Activity (2016).

[34]	R. J. Holt, JLAB Tritium Tech Note 1 (2012).

[35]	D. Meekins, Thermal Stress Analysis in Entrance Window (2015).

[36]	D. Meekins, Thermal Stress Analysis of Main Body Author (2015).

[37]	K. Welch, Thomas Jefferson National Accelerator Facility Jefferson Lab Process for Release of Material From Radiological Control (2010).

[38]	K. Welch, E. Abkemeier, and B. May, Technical Basis Document for Radioactivity Limits in Liquids as a Result of Activation or Contamination (2006).

[39]	P. Y. Pan and L. D. Rigdon, LANL (1996).

[40]	E. J. Beise, B. Brajuskovic, R. J. Holt, A. Ktrama, W. Korsch, T. O’Conner, G. G. Petratos, R. Ransome, P. Solvignon, and B. Wojtsekhowski, Tritium Gas Target Hazard Analysis for Jefferson Lab (2012).