## Slide 1: Radiological Readiness for Tritium Experiments
March 16, 2016
Keith Welch, MS, CHP
RadCon Deputy Manager

## Slide 2: Overview
Work-area and environmental monitoring
Personnel monitoring
Training
Operational controls
Target shipping/receipt/storage
Documentation
External communications/Public affairs

## Slide 3: Work area and environmental monitoring
Ambient hall air
  Two redundant systems
  At least one monitor must be operational when target is in hall
  Audible alarm, interlocked to exhaust, FSD
    Will base setpoint on HTO DAC
    Detection is ~0.005% of target escaping to the hall (over few hours or less)
Target chamber vacuum exhaust
  With 40 lpm purge, expect up to 1E-5 uCi/ml
  Operator alarm but not FSD
Response time < 20 seconds
Continuous air monitoring (CAM)

## Slide 4: Work area and environmental monitoring
Status and plans
  Use existing sample lines for ambient monitoring (retain existing monitoring of activation gases)
  Existing instrument checkout underway
    One channel ambient, one channel vac exhaust
  RFQ in on second instrument (~$20k)
  Need to do:
    New sample lines for vac exhaust (summer)
    Installation (fall)
    Interface to ventilation, EPICS, FSD (fall)
Continuous air monitoring (CAM)

## Slide 5: Work area and environmental monitoring
Continuous air monitoring (CAM)

## Slide 6: Work area and environmental monitoring
Portable air samplers used for receipt, during all target handling, and for any target chamber work
  Verbal agreement from SRS for loan
  Goal to have on site by October
Passive air sampling in adjacent spaces
  Diffusion vials (simple, inexpensive)
Hall will be routinely surveyed for H3 on surfaces
Routine monitoring of discharges, surface waters and groundwater
Equipment and procedures for contamination and environmental samples in place
Other Monitoring

## Slide 7: Personnel monitoring
Bioassay screening to be conducted
  Not a monitoring program
  Selected personnel from target group, hall workers, controls
  Expected number of workers ~ 10
  Occ Med to collect, ship samples
Status and plan
  Evaluating offerings by SRS and vendors
    Cost is comparable
    Plan to have baseline sampling in place ~October
Bioassay

## Slide 8: Training
Two-level training planned
  Level 1 for all personnel who can enter the hall
    General awareness training, response to alarms,
    Hall walkthrough (given by Hall A personnel)
  Level 2 for workers with access to “tritium zone”
    Focus on work practices, COO
  RadCon and target group to develop objectives
  RadCon will prepare and publish training
  Target date for completion October

## Slide 9: Operational Controls
Radiation Work Permits
  Established process, formal approvals, archived
  Target receipt sampling
  Work in hall
    Cover general hall access and target area work
  Can be tied to CANS system
Effluent controls
  Floor drain sump will be placed in manual mode
    Admin configuration control on pumps

## Slide 10: Target Shipping/Receipt/Storage
Target to be shipped in DOT Type B container (BTSP)
  Approved for 1.5 MCi
Trained SRS personnel conduct shipment both to and from JLab
DOE Orders and guides provide specific allowances for materials stored in certified Type B containers (exempt from emergency hazard assessments, nuclear facility inventories, etc.)
Container should be considered acceptable to meet all JLab containment requirements; design exceeds all local containment/confinement capabilities

## Slide 11: Target Shipping/Receipt/Storage
Receipt
  Receipt surveys required (per CFR)
  Maintain target in BTSP within hall upon receipt until ready to install
  Open BTSP in Hall remove secondary container to Hut
  Follow SRS procedures for sampling upon opening
Storage
  Use BTSP for all interim storage (weather emergencies, summer SAD, etc.)
  Consider using small “mover box” for enhanced security when in storage
  Maintain continuous air monitoring when target is in the hall

## Slide 12: Documentation
RadCon documentation process
  Work Permits
    RWPs are used to implement 10 CFR 835 requirements for work authorizations in radiological areas. They also satisfy work-control document requirements in the ES&H Manual. Online RWP system defines and requires management authorizations, provides a permanent, secure record.
    RWPs will be used to define scope of initial access procedures for ARMs and RadCon Techs (both of these procedures already exist), and to cover tasks involving receipt testing of target container and any work in the Hall involving the target chamber when H3 target is installed.
  Technical basis documents (TBDs)
    TBDs are used within the RadCon Department to define requirements for radiological program elements and document analyses used to justify rationale for a given technical approach. This is appropriate for documentation of such items as instrumentation configuration, assumptions used for emergency response, etc. TBDs are internally peer reviewed and approved by (at least) the RadCon Manager.
  Health Physics Implementing procedures (HPPs)
    HPPs cover routine tasks such as instrument calibration, lab analyses, surveys, etc.
All of these document types require review and authorization, have revision control and tracking and are permanently archived. All are audited as part of the DOE-approved Radiation Protection Program.

## Slide 13: Documentation
Status
TBDs | Status/target date
- Stack height to limit dose to the public | Complete
- Bioassay screening (rev existing TBD) | 7/30/16
- Accident/emergency response guidelines | 7/30/16
- Configuration of fixed air samplers | 5/31/16
- Deployment of passive monitors (rev existing TBD) | 7/30/16
Procedures | 
- Calibration and response checks of fixed monitors | 7/30/16
- Receipt of Type A quantity (existing doc) | Complete
- Operations of portable air monitors (SRS) | 9/30/16
- Leak-checking shipping container (SRS) | 9/30/16
Other | 
- Emergency Response Guidebook update | 9/30/16
- Public Affairs media package | 7/30/16
- Briefing for first responders | 7/30/16
- FSAD hazard analysis revision | 7/30/16
- RSAD (draft) | 9/30/16

## Slide 14: External Communications/Public Affairs
Strategy for public communication being developed in collaboration with Emerg. Mgmt and Pub Affairs
Two-tiered strategy adopted
  Proactive for first responders; brief NNFD on implications of accident (almost none), and modify online first responder’s guide
  Reactive for public in event of accident or media inquiry; prepare media packets and statements
  Documents under development, should be complete by summer