## Slide 1: Response to Target ERR Recommendations/Comments
ERR for Tritium Target 15 September 2015
Report was issued
  General statements
  Specific comments
  Specific recommendations
Summarized as Action Items
  Addressed in detail in Tritium Target Design Report

## Slide 2: General Action Items
Definition of Authorities
Definition of Procedures
Certification of calculations/designs
Hot checkout tests/simulations of equipment
Training/education
Emergency Planning
Public Awareness

## Slide 3: Authorities
Name | Group | Field
Dave Meekins | TGT | Tritium Target DA
Carrol Jones | FMM | HVAC/Smoke removal
Ed Douberly | FMFC | FPE/Smoke removal
Tom Renzo | FMF | Stack (PESE)
T. Michaelides | EESSAF | Machine FSD
Robin Wines | PHALLA | Reviewing DA/Hall A Eng
Tim Whitlatch | PHALLD | Reviewing DA
Keith Welch | RadCon | Radiological Control
Greg Howard | SRTE | Certification of Exhaust
Paul Blanton* | SRNL/SRTE | BTSP/Shipping
Ed Folts | PHY | DSO
Chris Keith | TGT | Target Processes
Jessie Butler | PHALLA | Hall A Work Coordinator (HAWC)
Jenord Alston | ENG | Examiner/Inspector
Mike Martin | QA/CI | CWI
John Warren | JLAB PA | Public Affairs

## Slide 4: Procedure Categories
Assembly, Testing, Examination and Inspection Procedures
  QC procedures for component acceptance
  Welding examination
  Leak/pressure testing
Hall access procedures
  Controlled, Truck Ramp Use, Restricted, etc.
Interlock Certification/checklist
FSD Certification/checklist
Exhaust System
  Hut flow/leak certification
  Interlock testing/certification
  Fire protection/high flow certification
Beam Operations Procedures
  Target Centering
  Beam Commissioning

## Slide 5: Procedures
OSP for target
  General target operation
  Specific critical tasks
    Installation and removal of cell
  Reference many task specific procedures
OSP/TOSP/AHA
  Standard signoff with appropriate SMEs
  Shall reference task specific procedures as applicable
Task specific procedures
  Reviewed/E-signed by authorities
  Electronically stored in JLAB Document Repository (same as official drawings)
  General target operation
Checkout procedures
  Used in JLAB Hot Checkout tool
  Multiple groups complete checklists and submit in HALOG/Target Log
  Final readiness approval by Hall A Lead.
RadCon Procedures
  RWPs and calibration procedures
  Bioassay procedure(s)
  Stored in RadCon document system
  Hall A access RWP

## Slide 6: Task Specific Procedure List
Number | Subject/Task | Approval
TGT-PROC-15-001 | Tritium target cell component cleaning | TGT DA
TGT-PROC-15-002 | Tritium target assembly and mechanical examination | TGT DA
TGT-PROC-15-003 | Tritium target cell leak testing procedure | TGT DA
TGT-PROC-15-004 | Tritium cell/BTSP receipt inspection | TGT,RadCon,FMSR
TGT-PROC-15-005 | Tritium target FSD testing and checkout | TGT,DSO,ACCEL
TGT-PROC-15-006 | Tritium target handling hut installation and removal | TGT,DSO
TGT-PROC-15-007 | Tritium target cell installation and removal | TGT,DSO
TGT-PROC-15-008 | Tritium target exhaust testing/certification | TGT,DSO,RadCon,IH,FML,FPE
TGT-PROC-15-009 | Tritium target truck ramp access | TGT,DSO,HAWC
TGT-PROC-15-010 | Tritium target general target operating procedure | TGT, DA
TGT-PROC-15-011 | General crane/material handling operations in tritium zone | TGT,RadCon,HAWC,DSO
TGT-PROC-15-012 | General maintenance work in tritium zone | TGT,RadCon,HAWC,DSO
TGT-PROC-15-013 | Tritium target collimator/seive slit installation/removal | TGT,RadCon,HAWC,DSO
TGT-PROC-15-014 | Tritium target Hall A prerun checklist | TGT,RadCon,HAWC,DSO

## Slide 7: RadCon Procedures
Addressed by Keith Welch

## Slide 8: General Procedures
Description | Approval
General tritium target OSP | DSO,SMEs
Tritium Zone Material handling TOSP(s) | DSO,SMEs,RadCon
FSD Checklist (Hot Checkout) | TGT,FSD,DSO
Interlock Checklist (Hot Checkout) | TGT,FSD,FML,DSO
WPS-015 | JLAB CWI
NDT-100 | QA/CI
ATS-RT procedure | DA, JLAB CWI

## Slide 9: Certification of Calculations/Design
Calculations
  Physics
  Radiological
  Pressure system
  Exhaust system Mechanical
  Exhaust system Structural
Design
  Vacuum
  Exhaust
  Pressure system

## Slide 10: Physics Calculations
Document Number | Description | Reviewed by
TGT-CALC-103-003 | Electron beam energy loss in cell | CDK
TGT-CALC-103-005 | Electron beam energy loss in Be isolation window | CDK
TGT-CALC-103-006 | Multiple scattering of beam in Be isolation window | CDK
TGT-CALC-103-009 | Thermal analysis of Be window in beam | CDK
TGT-CALC-103-010 | Tritium cell pressure and permeation rate | CDK
TGT-CALC-103-011 | Hot Spot simulations of tritium target release | KW
TGT-CALC-103-016 | CFD simulation of T2 gas in beam | N/A

## Slide 11: Pressure System Calculations
Document Number | Description | Reviewed by
TGT-CALC-103-002 | Pressure calculations for cell | TW/RW
TGT-CALC-103-007 | Elastic plastic analysis of cell assembly with shipping covers | TW/RW
TGT-CALC-103-008 | Thermal analysis of cell entrance window in beam | TW/RW
TGT-CALC-103-012 | Thermal stress analysis of the entrance window | N/A
TGT-CALC-103-013 | Thermal analysis of main body | TW/RW
TGT-CALC-103-014 | Thermal stress analysis of main body | N/A
TGT-CALC-103-015 | Thermal analysis of exit window with no raster | N/A
TGT-CALC-103-017 | Thermal/mech model of exit with no raster | N/A
TGT-CALC-103-018 | Thermal analysis of cell in ambient temperature | SRTE/SRNL

## Slide 12
Document Number | Description | Reviewed by
TGT-CALC-103-001 | Tritium exhaust ventilation | CJ, KW TBD
TGT-CALC-103-004 | Release of tritium into hall and worker exposure | KW TBD
TGT-CALC-103-006 | Multiple scattering of beam in Be isolation window | CDK
TGT-CALC-103-009 | Thermal analysis of Be window in beam | CDK
TGT-CALC-103-010 | Tritium cell pressure and permeation rate | CDK/BS et al. Tech Rep.
TGT-CALC-103-011 | Hot Spot simulations of tritium target release | KW TBD
Design
Exhaust/RadCon Calculations
Document Number | Description | Reviewed by
TGT-103-1001-0000 | System P&ID | DGM, CDK, CJ, KW, ED,
A11112-14-01-0800 | Stack top level | KW, DGM, CJ, ED, KW, TR
TGT-103-1000-0013 | Cell Assembly | DGM, RW, CDK
 | Exhaust system certification | Test
 | Vacuum system certification | Test

## Slide 13: Hot Checkout
Using Hot Checkout Tool
Similar to Hall D
  Tritium specific subsystems, components, and Check lists only
  Includes all sub systems affected by tritium
Performed prior to run start and after long down
Completed checklists shall be filed in HALOG/TARGETLOG with two signatures
Multiple levels of sign off at component, subsystem, and system levels.
Final Ready State set by Hall A lead

## Slide 14: Hot Checkout/Equipment Certification
Exhaust system
  Certified by testing/measurement (SRTE SME)
    Teach JLAB to repeat certification
    Fine tuning will also be set
  Calibrated equipment purchased
Vacuum system
  Standard leak test (HALOG/TARGETLOG logged)
Interlocks
  Issue fault condition for all interlocks and verify
FSD
  Issue fault condition for all FSD and verify
Verify CANS access

## Slide 16: Exhaust System Activation

## Slide 17: Vacuum Fault System
FSD issued on all faults
Getter operates on cold cathode failure
Pressure switch closes main turbo valves
Flow switch on water cooling for Be window
All faults trip alarms on UI
Pressure events trip heater power.

## Slide 18: Training Requirements
Standard Hall A training
Tritium I: General Hall A Access
  Web based RadCon training to be developed by M. Keller
  Special walkthrough of Hall A (Hall A personnel)
  Work in Hall A but outside Tritium Zone
  All Hall A shift workers/ARMS
Tritium II: Required for Tritium Zone work
  Classroom RadCon training for work inside Tritium Zone (while target is installed to be developed by M. Keller
Tritium target operator training: Analogous to Hall A cryotarget training.
BTSP: Training required for sealing and shipping BTSP
  Only required by SRS personnel performing these functions. JLAB personnel shall not perform these functions
  SRS personnel not required to take Tritium I and II.
Other training as required by JLAB ES&H Manual

## Slide 19: Emergency Planning
Emergency Response
  Local fire/EMS etc. (K. Welch)
Weather - Hurricane prep
  Warm up target
  Deenergize all power supplies
  Relieve vacuum
  Remove cell and store in BTSP/locked storage container

## Slide 20: Local Emergency Response
Emergency response governed by Emergency Responder’s Guidebook.
RadCon TBD to address minor edits
Only minor changes are expected
K. Welch is responsible for issue

## Slide 21: Public Affairs
NEPA: Supplemental analysis required
  File with DOE Headquarters
  DOE Function
Public Awareness
  JLAB Management
  JLAB PA

## Slide 22: Staffing levels
Normal staffing levels are expected for standard operations
Limited staff and controlled access for special conditions
  Target cell installation/removal
  Material handling inside Tritium Zone
  Unforeseen issues (all governed by TOSP/OSP)
  Staffing limited to personnel required to perform the work safely (RadCon, Hall A Tech Staff, Target Group).
  Staffing will be determined by TOSP/OSP and RWP
Access to Hall under special conditions shall be determined by Hall A work coordinator.

## Slide 23: Tritium Target Final Ladder
Target | Total Solid Thickness | Fluid Thickness
Tritium Cell | 0.16 g/cm2 | 0.075 g/cm2
Helium 3 Cell | 0.16 g/cm2 | 0.075 g/cm2
Deuterium Cell | 0.16 g/cm2 | 0.12 g/cm2
Empty Cell | 0.16 g/cm2 | N/A
Hydrogen Cell | 0.16 g/cm2 | 0.05 g/cm2
Aluminum Dummy | 0.3 g/cm2 | N/A
Carbon Optics | 0.2 g/cm2 | N/A
Carbon Foil | 0.167 g/cm2 | N/A
Carbon Hole | 0.167 g/cm2 | N/A
Thick Aluminum | 1.37g/cm2 | N/A
BeO | 0.16 g/cm2 | N/A
Be Isolation window | 0.04 g/cm2 | N/A

## Slide 24: Training Requirements
Two new training skills are required
  Tritium I: General access to Hall A for performing general work not near pivot/target
  Tritium II: Special training required only for individuals working near pivot/target when target is installed.
Training development lead: Maya Keller
Tritium I:
  Web based with tritium specific tour of Hall A
  Required by all those entering Hall A. Escort of untrained not allowed outside of emergency response
Tritium II:
  Special class for those individuals working neat pivot
  Target group, Hall A Tech Staff, others (working on chamber/beamline)

## Slide 25: Access Controls
CANS on ALL access points
  Tritium I is required for all personnel doors to be opened.
  All doors shall be posted with tritium signage.
Truck ramp doors controlled by CANS
  Limited to Hall A Technical Staff and Target Group Lead
  Tritium II is required to activate as well as being on limited list
No limits on staffing in Hall A are foreseen with exception of work near target/pivot and during critical operations
Controlled access protocols shall be in use during installation, removal, and other critical material handling operations on or near the target when it is in place.
  Access shall be limited to essential personnel
  Work shall be governed by OSP/TOSP and RWP as applicable

## Slide 26: Crane Use/Material Handling
Two cases:
  Lifts outside of Tritium Zone
  Lifts inside Tritium Zone
  Normal ESH requirements shall be met
Lifts/handling inside Zone:
  TOSP required
  Tritium II Training required
  Lift plan required
  RWP/RadCon approval of TOSP required
Lifts/handling outside zone:
  No special requirements
  Normal ES&H requirements
Truck ramp access has special procedure

## Slide 27: Admin Items
Task 1 Complete
    Hall A and Hut are considered “Confinement”, not “Containment/confinement” per DOE Handbook on Tritium Handling and Safe Storage.
  Documentation has been changed to be consistent with this definition.
Task 2 Working
    Safety controls must be considered in the Facility Safety Assessment Document (FSAD).
  The tritium target is considered an USI. This is addressed in the FSAD update. R. May is responsible for completing this task.
Task 3 Complete
    Explore using OTR from the Be window as a beam diagnostic.
  The possibility of using the Be window for an OTR has been investigated. The principle has advantages however due to close proximity of the window to the target face, the Be window is not visible from downstream. Unfortunately this window will also be difficult to see from the upstream side as well due to the narrow collimator installed. Therefore, using OTR on the Be window as a beam diagnostic has been determined to be not practical.

## Slide 28
Task 4 Working
    Develop/Implement public information campaign for employees and public (Lab Leadership) on science case and engineered safety controls. Have contextual information ready in case of incident. Publish article discussing tritium target issues.
  This process has been started. Upper Management, RadCon, and JLAB Public Affairs are responsible for completing this task.
Task 5 Working
    Educate NNFD on the hazards of Tritium so as to assist them in preparing to respond to an emergency event (Emergency Manager).
  Meeting is planned with NNFD. Minimal to no additional requirements for emergency response are anticipated. K. Welch is lead for addressing this issue.
Task 6 Complete
    Institute CANS for access control, evaluate for the Green Door.
  A CANS access point has been installed on this door. Additionally, another CANS access point is scheduled for installation at the top of the Hall A truck ramp to activate the rollup door. This access point shall only be available from the inside of the ramp. FML is completing the final installation in Spring 2016.

## Slide 29: Equipment Items
Task 1 Working
    We agree with “freezing” the design of the target and not allowing changes without an AD-level review/approval.
  Target design has been modified slightly. This was necessary for the cell assembly to fit into the BTSP. The alteration of the design affected the end shipping cover and the valve subassembly and did not affect the cell entrance window or main body.
  Physics AD approval of this change is pending the review of the design by R. Wines and T. Whitlatch.
Task 2 Complete
    We support the goal that the final target ladder configuration be finalized by March 1, 2016.
  The target ladder including gas targets, solid targets, and optics/dummy targets has been finalized.
Task 3 Complete
    No changes to the target ladder shall be allowed during the running of the four experiments.
  Agreed. No changes are planned or required for the ladder during the experiment.

## Slide 30
Task 4 Working
    All pressure system analyses provided during the review need to be reviewed and approved prior to February 1, 2016.
  Review of the pressure system analysis is currently underway. Alterations of the design required by shipping/packaging constraints resulted in a short delay in this process.
  Anticipated completion date for this review to be completed is 15 April 2016. R. Wines and T. Whitlatch are performing review.
  Review shall be documented using JLAB Document Repository.

## Slide 31
Task 5 Working
    It is important that detailed schematics of vacuum systems and exhaust systems presented today be checked or certified for functionality of what they are intended for. The as-built systems need to be simulated/tested before tritium target cell arrival from SRTE.
  The details of the vacuum and exhaust schematics are formally given on the system P&ID JLAB drawing TGT-103-1001-0000. This drawing is filed in the JLAB Document Repository. The following approvals (indicated by e-sign) shall be required:
    C. Jones FML
    JLAB Fire Protection Engineer (Ed Douberly)
    Target Group Lead C. Keith
    JLAB DA D. Meekins (Author)
    JLAB RadCon K. Welch
    Hall Engineering R. Wines
    Industrial Hygiene J. Williams
    Physics Division DSO E. Folts
  The vacuum system shall be certified by formal leak testing performed by a qualified JLAB vacuum leak test technician (P. Hood). The results of the leak test shall be documented and filed in the appropriate pressure system folder.
  The exhaust system shall be certified for two cases:
    Case 1: Low flow for Hut
    Case 2: High flow
    In both cases a report giving results of certifications shall be forwarded to FML, JLAB FPE, and RadCon. These reports shall be filed in the appropriate pressure system folder. SRTE Engineer will perform and teach JLAB DA to certify

## Slide 32
Task 6 Complete
    Consider adding unique identifiers (aka color/names) to the target gas cells.
  Each component shall be scribed with a unique JLAB serial number.
  Each cell assembly shall be examined using procedure TGT-PROC-15-002.
  Using the cell filling procedure (TGT-PROC-15-003), the cell fluid is recorded and permanently documented.
  This provides full traceability for the target cell and fill fluid. Note tritium cell assemblies shall be labeled JLAB-T2-01 , -02 etc. (Also required by SRTE)
Task 7 Complete
    Consider the collateral damage if the adjacent cell has a catastrophic failure.
  Note that the mechanical and chemical stored energy contained in each of the cells is extremely low.
  The only credible source of external heat affecting a single cell (e.g. not fire) is that of beam heating. Numerous safeguards exist that mitigate the risk of excessive beam heating
  The cell will fail in a ductile mode any debris from a failure would be directed out away from other cells in the stack which would remain unaffected.
  Experimental Run shall be stopped if cell failure is observed.
Task 8 Complete
    Consider physical keys on the target cells/ladder to ensure that each target cell can only be mounted at one location.
  Each target cell is uniquely labeled with full documented traceability. The use of unique keys would require at least one full set of spares. It is the opinion of the DA that such a standardization is the better choice in this case.

## Slide 33
Task 9 Complete
    Determine the maximum thickness and location for an exit window that does not negatively impact the physics program.
  It is clear that a Be window located about 50cm downstream from the target would negatively impact the physics from a data collection standpoint. This would however increase the dose for anyone working at or near the target/pivot. The relatively low tritium exposure to the exit pipe is unlikely to affect future work on this component due to current contamination. The cost of replacing this component is also not prohibitive. Tin summary, a 0.015 inch thick window would not negatively impact the physics but, would likely result in higher doses to personnel working near the pivot. Note that collimator and sieve slit changes are planned.
Task 10 Working
    Stress corrosion testing of pre-cracked Aluminum samples should be evaluated after 4, 8 and 12 months as planned and information should be utilized to determine the range of safe life-cycles of target cells.
  Testing is currently underway. Results are not yet available. These results shall be formalized and reported to Physics DSO.

## Slide 34: ES&H Items
Task 1 Working
    We agree with “freezing” the design of the target and not allowing changes without an AD-level review/approval.
  Target design has been modified slightly. This was necessary for the cell assembly to fit into the BTSP. The alteration of the design affected the end shipping cover and the valve subassembly and did not affect the cell entrance window or main body. Physics AD approval of this change is pending the review of the design by R. Wines and T. Whitlatch.
Task 2 Complete
    We support the goal that the final target ladder configuration be finalized by March 1, 2016.
  The target ladder including gas targets, solid targets, and optics/dummy targets has been finalized.
Task 3 Complete
    No changes to the target ladder shall be allowed during the running of the four experiments.
  Agreed. No changes are planned or required for the ladder during the experiment.
Task 4 Working
    All pressure system analyses provided during the review need to be reviewed and approved prior to February 1, 2016.
  Review of the pressure system analysis is currently underway. Alterations of the design required by shipping/packaging constraints resulted in a short delay in this process. Anticipated completion date for this review to be completed is 15 April 2016. R. Wines and T. Whitlatch are performing review. Review shall be documented using JLAB Document Repository

## Slide 35: Shipping Items
Task 1 Complete
    The transportation requirements for shipping the container must be fully understood and implemented with respect to flammable gas and radioactive material requirements.
  The tritium cell will contain a type A quantity of tritium inside a user specified container. The Bulk Tritium Shipping Package has been approved for use at SRS. The BTSP is certified to ship type B quantities contained in a user specified package. The BTSP is fully acceptable for use as the shipping container for the tritium cell.
  The BTSP exceeds the requirements for shipping hazardous material, type A quantities of Radioactive materials, and this quantity of flammable gas.
Task 2 Issues
    Arrange for just-in-time delivery of the cell to minimize the potential for outgassing within the shipping container.
  The complexity of scheduling, filling, leak testing, packaging and shipping at SRTE will not allow for just-in-time delivery of the tritium cell. Therefore two options exist. 1) JLAB intentionally delays filling/shipping until the run is expected to start. 2) The tritium cell is shipped near the run start date and will be received in the Hall. The cell will remain in the Hall inside the BTSP until it is installed.

## Slide 36
Task 3 Working
    Perform a dry-run with helium to validate design of packaging methodology. Consider dry-running forklift and installation operations as well.
  Each target cell shall undergo substantial leak testing which is summarized below:
  Leak test with He in reverse mode 500 psi internal pressure no covers installed.
  Leak test with He in reverse mode 1000 psi internal pressure with covers installed.
  Leak test on valve stem He in normal mode at 500 psi internal pressure
    No covers installed.
    Valve torqued to required setting
    Measures valve leak through
  Each leak test is documented and filed with pressure system documentation.
    Leak test is done on each cell.
    Traceable with serial numbers
  Note that there is no need to certify the BTSP which has undergone extensive testing and analysis. The MTV shall be secured inside the BTSP. Cell assembly mounted in MTV shall undergo drop testing at approved testing lab (Applied Testing Services).
  No fork lift dry run is planned.

## Slide 37
Task 4 Working
    Receiving should be done directly by the RadCon department. Inspection requirements will be defined with respect to acceptable limit of tritium contamination/outgassing.
  Receiving shall be performed by JLAB RadCon in collaboration with JLAB Target Group DA.
  The following procedure for receipt of the BTSP/cell is under development: TGT-PROC-16-004.
  JLAB RadCon and DA are working with SRNL to develop the procedure and determine acceptable limits for observed tritium outgassing.
Task 5 Working
    Keep shipping container staged nearby for potential storage needs. Evaluate need for alternate storage location.
  Agreed. The BTSP, when not containing the tritium cell, shall be stored in the Target Group Rad Area (note this is not a radiation area).
  The BTSP shall be used to store the cell when it is removed from the scattering chamber. The use of a locked and sealed storage box for the BTSP (containing the tritium cell) is being investigated. The storage location shall be

## Slide 38: Procedural Items
Task 1 Working
    The anticipated release dose relies on proper functioning of systems and correctly following of applicable procedures established for each release scenario. In the worst case scenario when no procedure is followed and all safety systems failed, a modest dose rate of 1.5 rem is anticipated. It is important that operational procedures are well established and followed by all personnel.
  Agreed. All personnel entering Hall A while the tritium target is located therein, shall receive training (Tritium I). Escorting personnel without this training shall not be allowed. This does not apply to SRS personnel involved in packaging/shipping the BTSP.
  Standardized procedures for general access and work in Hall A are under development.
  Detailed procedures for specific target tasks, beamline tasks and tasks near the pivot shall be numbered, rev’ed, and stored in the JLAB Document Repository. These procedures shall only be performed by trained and qualified personnel.
  Note that Tritium II training shall be required for all personnel working near the tritium cell.

## Slide 39
Task 2 Working
    The large number of administrative elements involved in maintaining the required beam size, beam current, target cell locations will require establishing detailed procedures and comprehensive training of relevant staff.
  Agreed. The following training shall be required for all personnel working in Hall A during the tritium run period.
  Tritium I: for general access and work in Hall A.
  Tritium II: for work near the tritium target or vacuum system while target is installed.
  Tritium I: General Hall A Access
    Web based RadCon training to be developed by M. Keller
    Work in Hall A but outside Tritium Zone
    All Hall A shift workers/ARMS
  Tritium II:
    Classroom RadCon training for work inside Tritium Zone to be developed by M. Keller
  Tritium target operator training: Analogous to Hall A cryotarget training.
  BTSP: Training required for sealing and shipping BTSP
    Only required by SRS personnel performing these functions. JLAB personnel shall not perform these functions
    SRS personnel not required to take Tritium I and II.
  Standard Hall A training
  Other training as required by JLAB ES&H Manual

## Slide 40
Task 3 Working
    Define restrictions/requirements during installation, changeover of equipment and post installation
      Staffing levels
      Access controls/CANS
      Training requirements
      Crane usage
      Local emergency response
      Et Cetera
    and present at the equipment-related readiness review.
  All issues addressed earlier in talk

## Slide 41
Task 4 Working
    Develop Hot Checkout Tool for Tritium Target with 2 layers of “hands-on” verification.
  A “Hot Checkout” list using a model similar to the Hall D Hot Checkout is under development. This system requires the following:
  Component level checklists performed by individuals
    Checklists shall be logged in HALOG/TargetLog
    Two signatures shall be required for checklist to be accepted.
  Subsystem level signoff by approved individuals
  System level signoff by approved individuals
  Final readiness signoff (by Hall leader or Physics DSO)
  Tool is under development by Hall A staff, et al.